## List1
Software will read settings from cells E2 to E6. |  |  | measurement properties |  |  | hardware data
Software will read measurement points from collumns A-C and |  |  | AC voltage amplitude | 1 | V | AC source idn | TEKTRONIX,AFG3101C,C010192,SCPI:99.0 FV:1.0.2
will write results into collumns D-H. Sheet must contain keywords |  |  | Sampling frequency | 15000000 | Hz | DC meter idn | HP3458A
START and STOP in collumn A, values of frequency in collumn B, |  |  | AC-DC settle time | 20 | s | Digitizer idn | NI PXI-5922, sn. F22D5B, channel 0
and record length in C. |  |  | AC-DC warm up time | 300 | s
|  |  | DC meter readings | 10
|  |  | Algorithm | TWM-WRMS
inputs |  |  | results
Command | Signal frequency | Record length | TIme of DC readings | Digitizer amplitude rms |  | DC voltage |  | Calculation |  | DC meter readings
(START, STOP) | Hz | multiple of periods | (time of writing this line into excel) | value (V) | unc (V) | value (V) | unc (V) | for figure |  | ¦---------------->
START | 10000 | 1000 | 2024-11-19 09:05 | 0.63785 | 0 | 0.54417 | 2.4384e-06 |  | 0.6383 | 0.54416 | 0.54416 | 0.54417 | 0.54417 | 0.54417
| 1000 | 1000 | 2024-11-19 09:06 | 0.63561 | 0 | 0.54379 | 0.00037752 | 1 | #N/A | 0.54417 | 0.54417 | 0.54416 | 0.54416 | 0.54228
| 20000 | 1000 | 2024-11-19 09:07 | 0.6364 | 0 | 0.54242 | 0.0001285 |  | 0.63596 | 0.54228 | 0.54229 | 0.54229 | 0.54229 | 0.54293
| 1000 | 1000 | 2024-11-19 09:08 | 0.6356 | 0 | 0.54279 | 0.00013234 | 1 | #N/A | 0.54292 | 0.54293 | 0.54292 | 0.54292 | 0.54226
| 30000 | 1000 | 2024-11-19 09:08 | 0.63584 | 0 | 0.54231 | 2.9058e-05 |  | 0.63575 | 0.54227 | 0.54228 | 0.54228 | 0.54228 | 0.54242
| 1000 | 1000 | 2024-11-19 09:09 | 0.63558 | 0 | 0.54238 | 3.2314e-05 | 1 | #N/A | 0.54242 | 0.54242 | 0.54241 | 0.54241 | 0.54225
| 40000 | 1000 | 2024-11-19 09:10 | 0.63572 | 0 | 0.54227 | 4.1728e-06 |  | 0.63571 | 0.54226 | 0.54226 | 0.54227 | 0.54227 | 0.54228
| 1000 | 1000 | 2024-11-19 09:10 | 0.6356 | 0 | 0.54228 | 3.7336e-06 | 1 | #N/A | 0.54228 | 0.54228 | 0.54228 | 0.54228 | 0.54226
| 50000 | 1000 | 2024-11-19 09:11 | 0.63564 | 0 | 0.54225 | 2.0116e-05 |  | 0.63573 | 0.54226 | 0.54227 | 0.54228 | 0.54228 | 0.54217
| 1000 | 1000 | 2024-11-19 09:12 | 0.63556 | 0 | 0.54217 | 1.5479e-05 | 1 | #N/A | 0.54217 | 0.54216 | 0.54215 | 0.54215 | 0.54223
| 60000 | 1000 | 2024-11-19 09:13 | 0.63566 | 0 | 0.54222 | 2.2188e-05 |  | 0.63574 | 0.54224 | 0.54224 | 0.54224 | 0.54225 | 0.54213
| 1000 | 1000 | 2024-11-19 09:13 | 0.63559 | 0 | 0.54215 | 2.7518e-05 | 1 | #N/A | 0.54213 | 0.54212 | 0.54212 | 0.54212 | 0.54226
| 70000 | 1001 | 2024-11-19 09:14 | 0.6357 | 0 | 0.54224 | 3.2933e-05 |  | 0.63584 | 0.54226 | 0.54226 | 0.54227 | 0.54228 | 0.5421
| 1000 | 1000 | 2024-11-19 09:15 | 0.63557 | 0 | 0.54212 | 3.1799e-05 | 1 | #N/A | 0.5421 | 0.54209 | 0.54208 | 0.54208 | 0.54224
| 80000 | 1000 | 2024-11-19 09:16 | 0.63578 | 0 | 0.54222 | 3.2599e-05 |  | 0.6359 | 0.54225 | 0.54225 | 0.54226 | 0.54227 | 0.54209
| 1000 | 1000 | 2024-11-19 09:16 | 0.63557 | 0 | 0.54212 | 3.1924e-05 | 1 | #N/A | 0.54209 | 0.54209 | 0.54208 | 0.54208 | 0.54224
| 90000 | 999 | 2024-11-19 09:17 | 0.63581 | 0 | 0.54222 | 4.4238e-05 |  | 0.63599 | 0.54225 | 0.54226 | 0.54227 | 0.54228 | 0.54204
| 1000 | 1000 | 2024-11-19 09:18 | 0.63557 | 0 | 0.54207 | 4.3782e-05 | 1 | #N/A | 0.54204 | 0.54203 | 0.54202 | 0.54201 | 0.54224
| 100000 | 1000 | 2024-11-19 09:19 | 0.63587 | 0 | 0.54221 | 5.0238e-05 |  | 0.63607 | 0.54225 | 0.54226 | 0.54226 | 0.54226 | 0.54201
| 1000 | 1000 | 2024-11-19 09:19 | 0.63555 | 0 | 0.54203 | 4.9244e-05 | 1 | #N/A | 0.542 | 0.54198 | 0.54198 | 0.54197 | 0.54223
| 110000 | 1001 | 2024-11-19 09:20 | 0.63597 | 0 | 0.5422 | 5.2399e-05 |  | 0.63617 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.54199
| 1000 | 1000 | 2024-11-19 09:21 | 0.63556 | 0 | 0.54202 | 5.2872e-05 | 1 | #N/A | 0.54198 | 0.54197 | 0.54197 | 0.54196 | 0.54223
| 120000 | 1000 | 2024-11-19 09:22 | 0.63608 | 0 | 0.54219 | 5.4149e-05 |  | 0.63628 | 0.54224 | 0.54225 | 0.54225 | 0.54225 | 0.54198
| 1000 | 1000 | 2024-11-19 09:22 | 0.63558 | 0 | 0.54202 | 5.7481e-05 | 1 | #N/A | 0.54198 | 0.54197 | 0.54196 | 0.54196 | 0.54225
| 130000 | 1001 | 2024-11-19 09:23 | 0.63617 | 0 | 0.5422 | 6.6943e-05 |  | 0.63643 | 0.54226 | 0.54226 | 0.54227 | 0.54228 | 0.54194
| 1000 | 1000 | 2024-11-19 09:24 | 0.63558 | 0 | 0.54198 | 6.7415e-05 | 1 | #N/A | 0.54193 | 0.54192 | 0.54191 | 0.5419 | 0.54225
| 140000 | 1001 | 2024-11-19 09:24 | 0.63629 | 0 | 0.5422 | 7.1132e-05 |  | 0.63657 | 0.54226 | 0.54227 | 0.54227 | 0.54227 | 0.54191
| 1000 | 1000 | 2024-11-19 09:25 | 0.63557 | 0 | 0.54196 | 6.9356e-05 | 1 | #N/A | 0.5419 | 0.54189 | 0.54189 | 0.54189 | 0.54224
| 150000 | 1000 | 2024-11-19 09:26 | 0.63638 | 0 | 0.54217 | 7.8974e-05 |  | 0.63668 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.54186
| 1000 | 1000 | 2024-11-19 09:27 | 0.63557 | 0 | 0.54192 | 8.045e-05 | 1 | #N/A | 0.54185 | 0.54184 | 0.54183 | 0.54183 | 0.54224
| 160000 | 1000 | 2024-11-19 09:27 | 0.63659 | 0 | 0.54219 | 7.2491e-05 |  | 0.63686 | 0.54224 | 0.54225 | 0.54227 | 0.54227 | 0.5419
| 1000 | 1000 | 2024-11-19 09:28 | 0.63558 | 0 | 0.54195 | 7.3882e-05 | 1 | #N/A | 0.54189 | 0.54188 | 0.54188 | 0.54187 | 0.54225
| 170000 | 1003 | 2024-11-19 09:29 | 0.63665 | 0 | 0.54217 | 9.1588e-05 |  | 0.637 | 0.54225 | 0.54226 | 0.54226 | 0.54226 | 0.5418
| 1000 | 1000 | 2024-11-19 09:30 | 0.63557 | 0 | 0.54187 | 9.2554e-05 | 1 | #N/A | 0.54179 | 0.54178 | 0.54177 | 0.54176 | 0.54224
| 180000 | 999 | 2024-11-19 09:30 | 0.63681 | 0 | 0.54216 | 9.2672e-05 |  | 0.63716 | 0.54225 | 0.54225 | 0.54226 | 0.54227 | 0.54179
| 1000 | 1000 | 2024-11-19 09:31 | 0.63557 | 0 | 0.54187 | 9.2789e-05 | 1 | #N/A | 0.54178 | 0.54178 | 0.54177 | 0.54176 | 0.54224
| 190000 | 1007 | 2024-11-19 09:32 | 0.63696 | 0 | 0.54215 | 9.8625e-05 |  | 0.63733 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.54176
| 1000 | 1000 | 2024-11-19 09:33 | 0.63556 | 0 | 0.54184 | 9.9328e-05 | 1 | #N/A | 0.54175 | 0.54174 | 0.54173 | 0.54173 | 0.54223
| 200000 | 1000 | 2024-11-19 09:33 | 0.63711 | 0 | 0.54215 | 0.00010611 |  | 0.6375 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.54172
| 1000 | 1000 | 2024-11-19 09:34 | 0.63556 | 0 | 0.54181 | 0.00010504 | 1 | #N/A | 0.54171 | 0.54171 | 0.5417 | 0.5417 | 0.54223
| 210000 | 1001 | 2024-11-19 09:35 | 0.63726 | 0 | 0.54213 | 0.00011164 |  | 0.63767 | 0.54224 | 0.54224 | 0.54225 | 0.54225 | 0.54169
| 1000 | 1000 | 2024-11-19 09:36 | 0.63556 | 0 | 0.54178 | 0.00011263 | 1 | #N/A | 0.54168 | 0.54167 | 0.54166 | 0.54166 | 0.54223
| 220000 | 1001 | 2024-11-19 09:36 | 0.63741 | 0 | 0.54213 | 0.0001206 |  | 0.63786 | 0.54224 | 0.54225 | 0.54225 | 0.54225 | 0.54165
| 1000 | 1000 | 2024-11-19 09:37 | 0.63554 | 0 | 0.54175 | 0.00011697 | 1 | #N/A | 0.54164 | 0.54163 | 0.54163 | 0.54162 | 0.54222
| 230400 | 1008 | 2024-11-19 09:38 | 0.63764 | 0 | 0.54212 | 0.00011639 |  | 0.63807 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.54166
| 1000 | 1000 | 2024-11-19 09:38 | 0.63557 | 0 | 0.54176 | 0.00011978 | 1 | #N/A | 0.54165 | 0.54164 | 0.54164 | 0.54163 | 0.54224
| 240000 | 1000 | 2024-11-19 09:39 | 0.6378 | 0 | 0.54212 | 0.00012516 |  | 0.63825 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.54162
| 1000 | 1000 | 2024-11-19 09:40 | 0.63558 | 0 | 0.54174 | 0.00012797 | 1 | #N/A | 0.54162 | 0.54161 | 0.5416 | 0.5416 | 0.54225
| 250000 | 1000 | 2024-11-19 09:41 | 0.63792 | 0 | 0.54211 | 0.00014364 |  | 0.63846 | 0.54225 | 0.54225 | 0.54226 | 0.54226 | 0.54154
| 1000 | 1000 | 2024-11-19 09:41 | 0.63555 | 0 | 0.54166 | 0.00014206 | 1 | #N/A | 0.54153 | 0.54152 | 0.54151 | 0.5415 | 0.54223
| 260000 | 1001 | 2024-11-19 09:42 | 0.6381 | 0 | 0.5421 | 0.000147 |  | 0.63865 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.54151
| 1000 | 1000 | 2024-11-19 09:43 | 0.63557 | 0 | 0.54164 | 0.00015133 | 1 | #N/A | 0.5415 | 0.54149 | 0.54148 | 0.54147 | 0.54224
| 270000 | 999 | 2024-11-19 09:44 | 0.63831 | 0 | 0.54211 | 0.00015199 |  | 0.63886 | 0.54225 | 0.54226 | 0.54226 | 0.54227 | 0.5415
| 1000 | 1000 | 2024-11-19 09:44 | 0.63559 | 0 | 0.54164 | 0.00015421 | 1 | #N/A | 0.54149 | 0.54149 | 0.54148 | 0.54147 | 0.54225
| 280000 | 1001 | 2024-11-19 09:45 | 0.63849 | 0 | 0.5421 | 0.00016098 |  | 0.63908 | 0.54226 | 0.54226 | 0.54227 | 0.54227 | 0.54146
| 1000 | 1000 | 2024-11-19 09:46 | 0.63558 | 0 | 0.54161 | 0.0001599 | 1 | #N/A | 0.54145 | 0.54145 | 0.54145 | 0.54145 | 0.54225
| 290000 | 986 | 2024-11-19 09:47 | 0.63864 | 0 | 0.54209 | 0.00017328 |  | 0.63927 | 0.54226 | 0.54226 | 0.54226 | 0.54226 | 0.54139
| 1000 | 1000 | 2024-11-19 09:47 | 0.63557 | 0 | 0.54156 | 0.00017168 | 1 | #N/A | 0.54139 | 0.54139 | 0.54138 | 0.54137 | 0.54224
| 300000 | 1000 | 2024-11-19 09:48 | 0.63882 | 0 | 0.54207 | 0.00018021 |  | 0.6395 | 0.54225 | 0.54225 | 0.54226 | 0.54226 | 0.54135
| 1000 | 1000 | 2024-11-19 09:49 | 0.63555 | 0 | 0.5415 | 0.00017928 | 1 | #N/A | 0.54134 | 0.54133 | 0.54132 | 0.54131 | 0.54222
| 310000 | 992 | 2024-11-19 09:50 | 0.63902 | 0 | 0.54206 | 0.00018177 |  | 0.63968 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.54133
| 1000 | 1000 | 2024-11-19 09:50 | 0.63557 | 0 | 0.5415 | 0.00018595 | 1 | #N/A | 0.54132 | 0.54132 | 0.54131 | 0.5413 | 0.54224
| 320000 | 1000 | 2024-11-19 09:51 | 0.63925 | 0 | 0.54207 | 0.00018246 |  | 0.63991 | 0.54225 | 0.54225 | 0.54225 | 0.54226 | 0.54134
| 1000 | 1000 | 2024-11-19 09:52 | 0.63557 | 0 | 0.54151 | 0.00018298 | 1 | #N/A | 0.54134 | 0.54134 | 0.54133 | 0.54132 | 0.54225
| 330000 | 1001 | 2024-11-19 09:52 | 0.63937 | 0 | 0.54206 | 0.00020531 |  | 0.64013 | 0.54226 | 0.54226 | 0.54227 | 0.54228 | 0.54124
| 1000 | 1000 | 2024-11-19 09:53 | 0.63556 | 0 | 0.54142 | 0.00020284 | 1 | #N/A | 0.54123 | 0.54122 | 0.54121 | 0.5412 | 0.54223
| 340000 | 1003 | 2024-11-19 09:54 | 0.63954 | 0 | 0.54204 | 0.00020933 |  | 0.64031 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.5412
| 1000 | 1000 | 2024-11-19 09:55 | 0.63555 | 0 | 0.54139 | 0.00020961 | 1 | #N/A | 0.54119 | 0.54118 | 0.54117 | 0.54117 | 0.54222
| 350000 | 1001 | 2024-11-19 09:55 | 0.63977 | 0 | 0.54203 | 0.00020685 |  | 0.64052 | 0.54223 | 0.54223 | 0.54224 | 0.54224 | 0.5412
| 1000 | 1000 | 2024-11-19 09:56 | 0.63556 | 0 | 0.5414 | 0.00020787 | 1 | #N/A | 0.5412 | 0.54119 | 0.54119 | 0.54118 | 0.54223
| 360000 | 999 | 2024-11-19 09:57 | 0.63995 | 0 | 0.54203 | 0.0002163 |  | 0.64074 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.54117
| 1000 | 1000 | 2024-11-19 09:58 | 0.63556 | 0 | 0.54137 | 0.00021543 | 1 | #N/A | 0.54116 | 0.54116 | 0.54115 | 0.54114 | 0.54223
| 370000 | 999 | 2024-11-19 09:58 | 0.64012 | 0 | 0.54203 | 0.0002238 |  | 0.64094 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.54113
| 1000 | 1000 | 2024-11-19 09:59 | 0.63555 | 0 | 0.54133 | 0.00022368 | 1 | #N/A | 0.54112 | 0.54111 | 0.54111 | 0.5411 | 0.54223
| 380000 | 1007 | 2024-11-19 10:00 | 0.64031 | 0 | 0.54202 | 0.00022686 |  | 0.64114 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.54111
| 1000 | 1000 | 2024-11-19 10:01 | 0.63556 | 0 | 0.54132 | 0.00022918 | 1 | #N/A | 0.5411 | 0.54109 | 0.54108 | 0.54107 | 0.54223
| 390000 | 1001 | 2024-11-19 10:01 | 0.64044 | 0 | 0.54201 | 0.00023902 |  | 0.6413 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.54105
| 1000 | 1000 | 2024-11-19 10:02 | 0.63556 | 0 | 0.54128 | 0.00023987 | 1 | #N/A | 0.54105 | 0.54104 | 0.54103 | 0.54103 | 0.54224
| 400000 | 1000 | 2024-11-19 10:03 | 0.64063 | 0 | 0.54201 | 0.00024344 |  | 0.64151 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.54103
| 1000 | 1000 | 2024-11-19 10:04 | 0.63556 | 0 | 0.54126 | 0.00024254 | 1 | #N/A | 0.54103 | 0.54102 | 0.54101 | 0.54101 | 0.54223
| 410000 | 984 | 2024-11-19 10:04 | 0.64079 | 0 | 0.542 | 0.00024899 |  | 0.64169 | 0.54224 | 0.54224 | 0.54225 | 0.54225 | 0.541
| 1000 | 1000 | 2024-11-19 10:05 | 0.63556 | 0 | 0.54124 | 0.00024871 | 1 | #N/A | 0.54099 | 0.541 | 0.54099 | 0.54098 | 0.54223
| 420000 | 1001 | 2024-11-19 10:06 | 0.64096 | 0 | 0.542 | 0.000253 |  | 0.64188 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.54098
| 1000 | 1000 | 2024-11-19 10:07 | 0.63558 | 0 | 0.54122 | 0.00025581 | 1 | #N/A | 0.54098 | 0.54097 | 0.54096 | 0.54096 | 0.54225
| 429000 | 1001 | 2024-11-19 10:07 | 0.64107 | 0 | 0.542 | 0.00026108 |  | 0.64202 | 0.54225 | 0.54226 | 0.54226 | 0.54227 | 0.54095
| 1000 | 1000 | 2024-11-19 10:08 | 0.63557 | 0 | 0.5412 | 0.00026138 | 1 | #N/A | 0.54094 | 0.54094 | 0.54093 | 0.54092 | 0.54224
| 440000 | 1001 | 2024-11-19 10:09 | 0.64124 | 0 | 0.54199 | 0.00026582 |  | 0.64221 | 0.54225 | 0.54225 | 0.54225 | 0.54226 | 0.54092
| 1000 | 1000 | 2024-11-19 10:09 | 0.63557 | 0 | 0.54117 | 0.00026774 | 1 | #N/A | 0.54092 | 0.5409 | 0.54089 | 0.54089 | 0.54224
| 450000 | 999 | 2024-11-19 10:10 | 0.64138 | 0 | 0.54199 | 0.00027313 |  | 0.64239 | 0.54226 | 0.54226 | 0.54227 | 0.54227 | 0.5409
| 1000 | 1000 | 2024-11-19 10:11 | 0.63555 | 0 | 0.54114 | 0.00027094 | 1 | #N/A | 0.54089 | 0.54088 | 0.54087 | 0.54086 | 0.54223
| 460000 | 989 | 2024-11-19 10:12 | 0.64151 | 0 | 0.54197 | 0.00027506 |  | 0.64251 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.54087
| 1000 | 1000 | 2024-11-19 10:12 | 0.63555 | 0 | 0.54112 | 0.00027475 | 1 | #N/A | 0.54086 | 0.54085 | 0.54085 | 0.54084 | 0.54222
| 470000 | 987 | 2024-11-19 10:13 | 0.64159 | 0 | 0.54195 | 0.00028666 |  | 0.64263 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.54081
| 1000 | 1000 | 2024-11-19 10:14 | 0.63555 | 0 | 0.54108 | 0.00028761 | 1 | #N/A | 0.5408 | 0.54079 | 0.54078 | 0.54078 | 0.54223
| 480000 | 1000 | 2024-11-19 10:15 | 0.64176 | 0 | 0.54197 | 0.00028424 |  | 0.64277 | 0.54224 | 0.54226 | 0.54226 | 0.54227 | 0.54084
| 1000 | 1000 | 2024-11-19 10:15 | 0.63559 | 0 | 0.54112 | 0.00028335 | 1 | #N/A | 0.54084 | 0.54084 | 0.54084 | 0.54084 | 0.54226
| 490500 | 981 | 2024-11-19 10:16 | 0.64187 | 0 | 0.54198 | 0.00029321 |  | 0.64293 | 0.54226 | 0.54227 | 0.54228 | 0.54228 | 0.54081
| 1000 | 1000 | 2024-11-19 10:17 | 0.63557 | 0 | 0.54108 | 0.00028923 | 1 | #N/A | 0.5408 | 0.5408 | 0.54079 | 0.54079 | 0.54224
| 500000 | 1000 | 2024-11-19 10:18 | 0.64194 | 0 | 0.54196 | 0.00029559 |  | 0.64302 | 0.54225 | 0.54225 | 0.54227 | 0.54228 | 0.54078
| 1000 | 1000 | 2024-11-19 10:18 | 0.63555 | 0 | 0.54105 | 0.00029311 | 1 | #N/A | 0.54077 | 0.54076 | 0.54076 | 0.54075 | 0.54223
| 510000 | 1003 | 2024-11-19 10:19 | 0.64203 | 0 | 0.54195 | 0.00029653 |  | 0.64312 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.54076
| 1000 | 1000 | 2024-11-19 10:20 | 0.63556 | 0 | 0.54103 | 0.00029953 | 1 | #N/A | 0.54075 | 0.54074 | 0.54073 | 0.54072 | 0.54223
| 520000 | 1001 | 2024-11-19 10:21 | 0.64212 | 0 | 0.54195 | 0.00029966 |  | 0.64321 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.54075
| 1000 | 1000 | 2024-11-19 10:21 | 0.63556 | 0 | 0.54103 | 0.00030162 | 1 | #N/A | 0.54074 | 0.54073 | 0.54072 | 0.54071 | 0.54223
| 530000 | 1007 | 2024-11-19 10:22 | 0.64214 | 0 | 0.54194 | 0.00031353 |  | 0.64328 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.54069
| 1000 | 1000 | 2024-11-19 10:23 | 0.63555 | 0 | 0.54098 | 0.00031058 | 1 | #N/A | 0.54068 | 0.54067 | 0.54066 | 0.54065 | 0.54222
| 540000 | 999 | 2024-11-19 10:23 | 0.64222 | 0 | 0.54193 | 0.0003079 |  | 0.64334 | 0.54223 | 0.54223 | 0.54224 | 0.54225 | 0.5407
| 1000 | 1000 | 2024-11-19 10:24 | 0.63555 | 0 | 0.54099 | 0.00030912 | 1 | #N/A | 0.54069 | 0.54068 | 0.54068 | 0.54067 | 0.54223
| 550000 | 1001 | 2024-11-19 10:25 | 0.64234 | 0 | 0.54194 | 0.00030057 |  | 0.64342 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.54074
| 1000 | 1000 | 2024-11-19 10:26 | 0.6356 | 0 | 0.54104 | 0.00030626 | 1 | #N/A | 0.54074 | 0.54073 | 0.54073 | 0.54073 | 0.54226
| 560000 | 994 | 2024-11-19 10:26 | 0.64232 | 0 | 0.54195 | 0.00031897 |  | 0.64348 | 0.54227 | 0.54227 | 0.54228 | 0.54228 | 0.54068
| 1000 | 1000 | 2024-11-19 10:27 | 0.63558 | 0 | 0.54097 | 0.00031918 | 1 | #N/A | 0.54067 | 0.54066 | 0.54065 | 0.54064 | 0.54225
| 570000 | 1007 | 2024-11-19 10:28 | 0.64229 | 0 | 0.54194 | 0.00032836 |  | 0.64349 | 0.54226 | 0.54227 | 0.54227 | 0.54227 | 0.54063
| 1000 | 1000 | 2024-11-19 10:29 | 0.63558 | 0 | 0.54093 | 0.00032928 | 1 | #N/A | 0.54062 | 0.54061 | 0.5406 | 0.54059 | 0.54225
| 580000 | 986 | 2024-11-19 10:29 | 0.64228 | 0 | 0.54193 | 0.00033546 |  | 0.64351 | 0.54226 | 0.54227 | 0.54227 | 0.54228 | 0.54059
| 1000 | 1000 | 2024-11-19 10:30 | 0.63557 | 0 | 0.5409 | 0.00033369 | 1 | #N/A | 0.54058 | 0.54057 | 0.54056 | 0.54056 | 0.54224
| 590000 | 1003 | 2024-11-19 10:31 | 0.64228 | 0 | 0.54192 | 0.00033434 |  | 0.64349 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.54058
| 1000 | 1000 | 2024-11-19 10:32 | 0.63556 | 0 | 0.5409 | 0.00033317 | 1 | #N/A | 0.54057 | 0.54056 | 0.54056 | 0.54055 | 0.54223
| 600000 | 1000 | 2024-11-19 10:32 | 0.64223 | 0 | 0.54191 | 0.00034056 |  | 0.64347 | 0.54223 | 0.54225 | 0.54226 | 0.54227 | 0.54055
| 1000 | 1000 | 2024-11-19 10:33 | 0.63556 | 0 | 0.54087 | 0.00034052 | 1 | #N/A | 0.54054 | 0.54053 | 0.54052 | 0.54052 | 0.54223
| 609600 | 1016 | 2024-11-19 10:34 | 0.64227 | 0 | 0.54192 | 0.00033383 |  | 0.64348 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.54058
| 1000 | 1000 | 2024-11-19 10:35 | 0.63558 | 0 | 0.5409 | 0.00033814 | 1 | #N/A | 0.54057 | 0.54056 | 0.54056 | 0.54055 | 0.54225
| 620000 | 992 | 2024-11-19 10:35 | 0.64216 | 0 | 0.54192 | 0.00034784 |  | 0.64342 | 0.54226 | 0.54226 | 0.54227 | 0.54228 | 0.54053
| 1000 | 1000 | 2024-11-19 10:36 | 0.63556 | 0 | 0.54085 | 0.00034334 | 1 | #N/A | 0.54052 | 0.54051 | 0.54051 | 0.5405 | 0.54223
| 630000 | 1008 | 2024-11-19 10:37 | 0.64207 | 0 | 0.5419 | 0.00035394 |  | 0.64336 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.54048
| 1000 | 1000 | 2024-11-19 10:38 | 0.63555 | 0 | 0.54081 | 0.00035247 | 1 | #N/A | 0.54047 | 0.54046 | 0.54046 | 0.54044 | 0.54222
| 640000 | 992 | 2024-11-19 10:38 | 0.64207 | 0 | 0.5419 | 0.00034154 |  | 0.6433 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.54053
| 1000 | 1000 | 2024-11-19 10:39 | 0.63558 | 0 | 0.54086 | 0.00034706 | 1 | #N/A | 0.54052 | 0.54051 | 0.54051 | 0.5405 | 0.54225
| 650000 | 1001 | 2024-11-19 10:40 | 0.6419 | 0 | 0.5419 | 0.0003652 |  | 0.64324 | 0.54226 | 0.54226 | 0.54227 | 0.54228 | 0.54044
| 1000 | 1000 | 2024-11-19 10:41 | 0.63555 | 0 | 0.54078 | 0.00036098 | 1 | #N/A | 0.54043 | 0.54042 | 0.54041 | 0.54041 | 0.54222
| 660000 | 1001 | 2024-11-19 10:41 | 0.64182 | 0 | 0.5419 | 0.00036359 |  | 0.64314 | 0.54223 | 0.54226 | 0.54227 | 0.54229 | 0.54044
| 1000 | 1000 | 2024-11-19 10:42 | 0.63557 | 0 | 0.54079 | 0.00036226 | 1 | #N/A | 0.54044 | 0.54043 | 0.54042 | 0.54041 | 0.54224
| 670000 | 1005 | 2024-11-19 10:43 | 0.64172 | 0 | 0.5419 | 0.00036376 |  | 0.64304 | 0.54225 | 0.54226 | 0.54226 | 0.54227 | 0.54044
| 1000 | 1000 | 2024-11-19 10:43 | 0.63557 | 0 | 0.54079 | 0.0003633 | 1 | #N/A | 0.54043 | 0.54042 | 0.54042 | 0.54042 | 0.54224
| 680000 | 1003 | 2024-11-19 10:44 | 0.64157 | 0 | 0.54188 | 0.00037126 |  | 0.64292 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.5404
| 1000 | 1000 | 2024-11-19 10:45 | 0.63556 | 0 | 0.54074 | 0.00037178 | 1 | #N/A | 0.54039 | 0.54038 | 0.54037 | 0.54036 | 0.54223
| 690000 | 989 | 2024-11-19 10:46 | 0.64141 | 0 | 0.54188 | 0.00037407 |  | 0.64277 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.54038
| 1000 | 1000 | 2024-11-19 10:46 | 0.63555 | 0 | 0.54073 | 0.00037392 | 1 | #N/A | 0.54037 | 0.54036 | 0.54036 | 0.54035 | 0.54223
| 700000 | 1001 | 2024-11-19 10:47 | 0.64128 | 0 | 0.54188 | 0.00037145 |  | 0.64263 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.54039
| 1000 | 1000 | 2024-11-19 10:48 | 0.63558 | 0 | 0.54074 | 0.00037484 | 1 | #N/A | 0.54038 | 0.54037 | 0.54036 | 0.54036 | 0.54224
| 710000 | 994 | 2024-11-19 10:49 | 0.64107 | 0 | 0.54188 | 0.00038674 |  | 0.64247 | 0.54225 | 0.54226 | 0.54227 | 0.54228 | 0.54033
| 1000 | 1000 | 2024-11-19 10:49 | 0.63558 | 0 | 0.54069 | 0.00038801 | 1 | #N/A | 0.54031 | 0.5403 | 0.5403 | 0.5403 | 0.54224
| 720000 | 1002 | 2024-11-19 10:50 | 0.6409 | 0 | 0.54187 | 0.00038717 |  | 0.64229 | 0.54225 | 0.54225 | 0.54226 | 0.54227 | 0.54032
| 1000 | 1000 | 2024-11-19 10:51 | 0.63557 | 0 | 0.54069 | 0.00038584 | 1 | #N/A | 0.54031 | 0.54031 | 0.54031 | 0.54031 | 0.54224
| 728000 | 1001 | 2024-11-19 10:52 | 0.64076 | 0 | 0.54187 | 0.00038292 |  | 0.64214 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.54034
| 1000 | 1000 | 2024-11-19 10:52 | 0.63556 | 0 | 0.54071 | 0.00038223 | 1 | #N/A | 0.54033 | 0.54033 | 0.54032 | 0.54031 | 0.54224
| 740000 | 999 | 2024-11-19 10:53 | 0.64048 | 0 | 0.54186 | 0.00039523 |  | 0.64192 | 0.54225 | 0.54225 | 0.54226 | 0.54227 | 0.54028
| 1000 | 1000 | 2024-11-19 10:54 | 0.63556 | 0 | 0.54065 | 0.00039623 | 1 | #N/A | 0.54027 | 0.54026 | 0.54025 | 0.54024 | 0.54223
| 750000 | 1000 | 2024-11-19 10:55 | 0.64025 | 0 | 0.54185 | 0.00039676 |  | 0.64168 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.54026
| 1000 | 1000 | 2024-11-19 10:55 | 0.63557 | 0 | 0.54064 | 0.00039824 | 1 | #N/A | 0.54025 | 0.54025 | 0.54024 | 0.54024 | 0.54224
| 760000 | 1007 | 2024-11-19 10:56 | 0.64002 | 0 | 0.54186 | 0.00039742 |  | 0.64145 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.54027
| 1000 | 1000 | 2024-11-19 10:57 | 0.63559 | 0 | 0.54065 | 0.00040099 | 1 | #N/A | 0.54026 | 0.54025 | 0.54025 | 0.54024 | 0.54226
| 770000 | 1001 | 2024-11-19 10:57 | 0.63974 | 0 | 0.54186 | 0.00041286 |  | 0.64124 | 0.54226 | 0.54227 | 0.54227 | 0.54228 | 0.54021
| 1000 | 1000 | 2024-11-19 10:58 | 0.63553 | 0 | 0.54058 | 0.00040717 | 1 | #N/A | 0.54019 | 0.54018 | 0.54017 | 0.54016 | 0.54221
| 780000 | 1001 | 2024-11-19 10:59 | 0.63954 | 0 | 0.54183 | 0.00040171 |  | 0.64099 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.54022
| 1000 | 1000 | 2024-11-19 11:00 | 0.63555 | 0 | 0.54061 | 0.00040341 | 1 | #N/A | 0.54022 | 0.54021 | 0.5402 | 0.54019 | 0.54222
| 790400 | 988 | 2024-11-19 11:00 | 0.63928 | 0 | 0.54184 | 0.00039991 |  | 0.64071 | 0.54223 | 0.54224 | 0.54224 | 0.54226 | 0.54024
| 1000 | 1000 | 2024-11-19 11:01 | 0.63558 | 0 | 0.54063 | 0.00040258 | 1 | #N/A | 0.54024 | 0.54023 | 0.54023 | 0.54022 | 0.54224
| 800000 | 1000 | 2024-11-19 11:02 | 0.63899 | 0 | 0.54184 | 0.00041227 |  | 0.64047 | 0.54225 | 0.54226 | 0.54226 | 0.54226 | 0.5402
| 1000 | 1000 | 2024-11-19 11:03 | 0.63556 | 0 | 0.54059 | 0.00040878 | 1 | #N/A | 0.54019 | 0.54019 | 0.54018 | 0.54017 | 0.54223
| 810000 | 999 | 2024-11-19 11:03 | 0.63867 | 0 | 0.54183 | 0.00041737 |  | 0.64017 | 0.54224 | 0.54225 | 0.54225 | 0.54227 | 0.54016
| 1000 | 1000 | 2024-11-19 11:04 | 0.63556 | 0 | 0.54056 | 0.00041689 | 1 | #N/A | 0.54015 | 0.54015 | 0.54014 | 0.54014 | 0.54223
| 820000 | 984 | 2024-11-19 11:05 | 0.63839 | 0 | 0.54183 | 0.0004171 |  | 0.63989 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.54016
| 1000 | 1000 | 2024-11-19 11:06 | 0.63555 | 0 | 0.54056 | 0.00041717 | 1 | #N/A | 0.54016 | 0.54014 | 0.54013 | 0.54013 | 0.54223
| 830000 | 996 | 2024-11-19 11:06 | 0.63812 | 0 | 0.54183 | 0.00041674 |  | 0.63961 | 0.54223 | 0.54224 | 0.54226 | 0.54226 | 0.54017
| 1000 | 1000 | 2024-11-19 11:07 | 0.63557 | 0 | 0.54056 | 0.00041899 | 1 | #N/A | 0.54015 | 0.54014 | 0.54014 | 0.54014 | 0.54224
| 840000 | 1001 | 2024-11-19 11:08 | 0.63776 | 0 | 0.54183 | 0.00042635 |  | 0.6393 | 0.54225 | 0.54225 | 0.54225 | 0.54226 | 0.54012
| 1000 | 1000 | 2024-11-19 11:09 | 0.63557 | 0 | 0.54052 | 0.00042809 | 1 | #N/A | 0.54011 | 0.5401 | 0.54009 | 0.54008 | 0.54223
| 850000 | 1003 | 2024-11-19 11:09 | 0.63742 | 0 | 0.54182 | 0.00043128 |  | 0.63898 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.54009
| 1000 | 1000 | 2024-11-19 11:10 | 0.63555 | 0 | 0.5405 | 0.0004313 | 1 | #N/A | 0.54008 | 0.54007 | 0.54006 | 0.54006 | 0.54222
| 860000 | 989 | 2024-11-19 11:11 | 0.63712 | 0 | 0.54182 | 0.00042531 |  | 0.63864 | 0.54223 | 0.54224 | 0.54224 | 0.54226 | 0.54012
| 1000 | 1000 | 2024-11-19 11:12 | 0.63558 | 0 | 0.54053 | 0.00042989 | 1 | #N/A | 0.54011 | 0.54011 | 0.5401 | 0.54009 | 0.54225
| 870000 | 986 | 2024-11-19 11:12 | 0.63676 | 0 | 0.54183 | 0.00043695 |  | 0.63834 | 0.54226 | 0.54226 | 0.54227 | 0.54227 | 0.54008
| 1000 | 1000 | 2024-11-19 11:13 | 0.63555 | 0 | 0.54048 | 0.00043354 | 1 | #N/A | 0.54007 | 0.54006 | 0.54004 | 0.54004 | 0.54222
| 880000 | 990 | 2024-11-19 11:14 | 0.63639 | 0 | 0.5418 | 0.00044214 |  | 0.63798 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.54003
| 1000 | 1000 | 2024-11-19 11:15 | 0.63555 | 0 | 0.54046 | 0.00044318 | 1 | #N/A | 0.54002 | 0.54002 | 0.54001 | 0.54 | 0.54223
| 888000 | 999 | 2024-11-19 11:15 | 0.63612 | 0 | 0.54181 | 0.00044277 |  | 0.63772 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.54004
| 1000 | 1000 | 2024-11-19 11:16 | 0.63555 | 0 | 0.54045 | 0.00044444 | 1 | #N/A | 0.54002 | 0.54001 | 0.54 | 0.53999 | 0.54223
| 900000 | 999 | 2024-11-19 11:17 | 0.63569 | 0 | 0.54181 | 0.00044889 |  | 0.6373 | 0.54224 | 0.54226 | 0.54226 | 0.54227 | 0.54001
| 1000 | 1000 | 2024-11-19 11:18 | 0.63555 | 0 | 0.54044 | 0.00044734 | 1 | #N/A | 0.54 | 0.53999 | 0.53998 | 0.53998 | 0.54223
| 910000 | 1001 | 2024-11-19 11:18 | 0.63537 | 0 | 0.54181 | 0.0004458 |  | 0.63697 | 0.54224 | 0.54226 | 0.54226 | 0.54227 | 0.54003
| 1000 | 1000 | 2024-11-19 11:19 | 0.63557 | 0 | 0.54045 | 0.00044647 | 1 | #N/A | 0.54002 | 0.54001 | 0.54 | 0.53999 | 0.54224
| 920000 | 989 | 2024-11-19 11:20 | 0.63494 | 0 | 0.5418 | 0.00045939 |  | 0.63658 | 0.54225 | 0.54226 | 0.54227 | 0.54228 | 0.53997
| 1000 | 1000 | 2024-11-19 11:20 | 0.63557 | 0 | 0.5404 | 0.00045824 | 1 | #N/A | 0.53996 | 0.53995 | 0.53994 | 0.53994 | 0.54224
| 930000 | 992 | 2024-11-19 11:21 | 0.63456 | 0 | 0.54179 | 0.00046227 |  | 0.63622 | 0.54225 | 0.54225 | 0.54226 | 0.54226 | 0.53994
| 1000 | 1000 | 2024-11-19 11:22 | 0.63554 | 0 | 0.54038 | 0.00045793 | 1 | #N/A | 0.53994 | 0.53993 | 0.53992 | 0.53991 | 0.54222
| 940000 | 987 | 2024-11-19 11:23 | 0.63418 | 0 | 0.54177 | 0.0004639 |  | 0.63583 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53992
| 1000 | 1000 | 2024-11-19 11:23 | 0.63556 | 0 | 0.54036 | 0.00046786 | 1 | #N/A | 0.53991 | 0.5399 | 0.53989 | 0.53989 | 0.54223
| 950000 | 1007 | 2024-11-19 11:24 | 0.63382 | 0 | 0.54178 | 0.00046805 |  | 0.63549 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.53991
| 1000 | 1000 | 2024-11-19 11:25 | 0.63555 | 0 | 0.54035 | 0.0004688 | 1 | #N/A | 0.5399 | 0.53989 | 0.53988 | 0.53987 | 0.54223
| 960000 | 1000 | 2024-11-19 11:26 | 0.6335 | 0 | 0.54178 | 0.00046027 |  | 0.63512 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.53994
| 1000 | 1000 | 2024-11-19 11:26 | 0.6356 | 0 | 0.54039 | 0.00046952 | 1 | #N/A | 0.53993 | 0.53992 | 0.53992 | 0.53992 | 0.54227
| 969600 | 1010 | 2024-11-19 11:27 | 0.63314 | 0 | 0.54181 | 0.00046807 |  | 0.63482 | 0.54227 | 0.54228 | 0.54228 | 0.54228 | 0.53994
| 1000 | 1000 | 2024-11-19 11:28 | 0.63558 | 0 | 0.54038 | 0.00046727 | 1 | #N/A | 0.53993 | 0.53992 | 0.53991 | 0.5399 | 0.54225
| 980000 | 980 | 2024-11-19 11:29 | 0.6327 | 0 | 0.54178 | 0.00047798 |  | 0.63441 | 0.54225 | 0.54226 | 0.54226 | 0.54227 | 0.53987
| 1000 | 1000 | 2024-11-19 11:29 | 0.63555 | 0 | 0.54032 | 0.0004763 | 1 | #N/A | 0.53986 | 0.53985 | 0.53984 | 0.53984 | 0.54223
| 990000 | 990 | 2024-11-19 11:30 | 0.63229 | 0 | 0.54177 | 0.00048355 |  | 0.63401 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.53983
| 1000 | 1000 | 2024-11-19 11:31 | 0.63557 | 0 | 0.5403 | 0.0004846 | 1 | #N/A | 0.53982 | 0.53981 | 0.53981 | 0.53981 | 0.54224
| 1000000 | 1000 | 2024-11-19 11:31 | 0.63197 | 0 | 0.54178 | 0.00047988 |  | 0.63368 | 0.54225 | 0.54225 | 0.54226 | 0.54226 | 0.53986
| 1000 | 1000 | 2024-11-19 11:32 | 0.63556 | 0 | 0.54031 | 0.00048128 | 1 | #N/A | 0.53985 | 0.53983 | 0.53982 | 0.53981 | 0.54223
| 1005000 | 1005 | 2024-11-19 11:33 | 0.63179 | 0 | 0.54177 | 0.0004819 |  | 0.63351 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.53984
| 1000 | 1000 | 2024-11-19 11:34 | 0.63556 | 0 | 0.5403 | 0.00048267 | 1 | #N/A | 0.53983 | 0.53983 | 0.53982 | 0.53981 | 0.54224
| 1020000 | 1003 | 2024-11-19 11:34 | 0.63121 | 0 | 0.54176 | 0.00049189 |  | 0.63296 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.53979
| 1000 | 1000 | 2024-11-19 11:35 | 0.63555 | 0 | 0.54026 | 0.00049189 | 1 | #N/A | 0.53978 | 0.53977 | 0.53976 | 0.53975 | 0.54222
| 1020000 | 1003 | 2024-11-19 11:36 | 0.63118 | 0 | 0.54176 | 0.00049399 |  | 0.63295 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.53978
| 1000 | 1000 | 2024-11-19 11:37 | 0.63556 | 0 | 0.54025 | 0.00049573 | 1 | #N/A | 0.53977 | 0.53976 | 0.53975 | 0.53974 | 0.54223
| 1050000 | 1001 | 2024-11-19 11:37 | 0.63016 | 0 | 0.54176 | 0.00049653 |  | 0.63193 | 0.54224 | 0.54225 | 0.54227 | 0.54227 | 0.53978
| 1000 | 1000 | 2024-11-19 11:38 | 0.63557 | 0 | 0.54025 | 0.00049684 | 1 | #N/A | 0.53976 | 0.53975 | 0.53974 | 0.53974 | 0.54223
| 1050000 | 1001 | 2024-11-19 11:39 | 0.63014 | 0 | 0.54176 | 0.00049731 |  | 0.6319 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.53977
| 1000 | 1000 | 2024-11-19 11:40 | 0.63557 | 0 | 0.54025 | 0.00049817 | 1 | #N/A | 0.53976 | 0.53975 | 0.53975 | 0.53973 | 0.54224
| 1060000 | 1007 | 2024-11-19 11:40 | 0.62977 | 0 | 0.54175 | 0.00050596 |  | 0.63157 | 0.54225 | 0.54225 | 0.54226 | 0.54227 | 0.53973
| 1000 | 1000 | 2024-11-19 11:41 | 0.63556 | 0 | 0.54021 | 0.00050544 | 1 | #N/A | 0.53972 | 0.53971 | 0.5397 | 0.53969 | 0.54223
| 1065000 | 994 | 2024-11-19 11:42 | 0.62961 | 0 | 0.54174 | 0.00050587 |  | 0.63139 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.53972
| 1000 | 1000 | 2024-11-19 11:43 | 0.63556 | 0 | 0.54021 | 0.00050567 | 1 | #N/A | 0.53972 | 0.53971 | 0.5397 | 0.5397 | 0.54223
| 1080000 | 999 | 2024-11-19 11:43 | 0.62913 | 0 | 0.54175 | 0.00050388 |  | 0.6309 | 0.54224 | 0.54224 | 0.54226 | 0.54226 | 0.53973
| 1000 | 1000 | 2024-11-19 11:44 | 0.63559 | 0 | 0.54023 | 0.00050847 | 1 | #N/A | 0.53972 | 0.53972 | 0.53972 | 0.53972 | 0.54226
| 1092000 | 1001 | 2024-11-19 11:45 | 0.62875 | 0 | 0.54176 | 0.0005123 |  | 0.63056 | 0.54227 | 0.54227 | 0.54227 | 0.54228 | 0.53971
| 1000 | 1000 | 2024-11-19 11:46 | 0.6356 | 0 | 0.54021 | 0.0005136 | 1 | #N/A | 0.5397 | 0.53969 | 0.53969 | 0.53969 | 0.54226
| 1100000 | 1001 | 2024-11-19 11:46 | 0.62855 | 0 | 0.54177 | 0.00050562 |  | 0.63031 | 0.54226 | 0.54227 | 0.54228 | 0.54229 | 0.53975
| 1000 | 1000 | 2024-11-19 11:47 | 0.63562 | 0 | 0.54025 | 0.00050708 | 1 | #N/A | 0.53975 | 0.53975 | 0.53974 | 0.53973 | 0.54228
| 1110000 | 999 | 2024-11-19 11:48 | 0.62818 | 0 | 0.54177 | 0.00052201 |  | 0.63004 | 0.54228 | 0.54229 | 0.54229 | 0.5423 | 0.53968
| 1000 | 1000 | 2024-11-19 11:48 | 0.63559 | 0 | 0.54017 | 0.00052267 | 1 | #N/A | 0.53967 | 0.53966 | 0.53964 | 0.53963 | 0.54226
| 1125000 | 999 | 2024-11-19 11:49 | 0.62767 | 0 | 0.54175 | 0.00053523 |  | 0.62958 | 0.54227 | 0.54228 | 0.54228 | 0.54229 | 0.5396
| 1000 | 1000 | 2024-11-19 11:50 | 0.63556 | 0 | 0.54011 | 0.0005307 | 1 | #N/A | 0.53959 | 0.53958 | 0.53957 | 0.53957 | 0.54223
| 1125000 | 999 | 2024-11-19 11:51 | 0.6277 | 0 | 0.54173 | 0.00052604 |  | 0.62956 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.53962
| 1000 | 1000 | 2024-11-19 11:51 | 0.63558 | 0 | 0.54013 | 0.00052984 | 1 | #N/A | 0.53962 | 0.5396 | 0.53959 | 0.53958 | 0.54225
| 1144000 | 1001 | 2024-11-19 11:52 | 0.62715 | 0 | 0.54173 | 0.00053803 |  | 0.62903 | 0.54226 | 0.54226 | 0.54227 | 0.54227 | 0.53957
| 1000 | 1000 | 2024-11-19 11:53 | 0.63558 | 0 | 0.5401 | 0.00053636 | 1 | #N/A | 0.53958 | 0.53957 | 0.53956 | 0.53955 | 0.54224
| 1155000 | 1001 | 2024-11-19 11:54 | 0.62684 | 0 | 0.54172 | 0.00054222 |  | 0.62877 | 0.54225 | 0.54226 | 0.54226 | 0.54227 | 0.53955
| 1000 | 1000 | 2024-11-19 11:54 | 0.63554 | 0 | 0.54006 | 0.00053861 | 1 | #N/A | 0.53955 | 0.53953 | 0.53952 | 0.53951 | 0.54222
| 1155000 | 1001 | 2024-11-19 11:55 | 0.62682 | 0 | 0.5417 | 0.00054285 |  | 0.62873 | 0.54223 | 0.54224 | 0.54224 | 0.54226 | 0.53953
| 1000 | 1000 | 2024-11-19 11:56 | 0.63554 | 0 | 0.54005 | 0.0005408 | 1 | #N/A | 0.53952 | 0.53951 | 0.53951 | 0.53951 | 0.54222
| 1180000 | 1003 | 2024-11-19 11:57 | 0.62626 | 0 | 0.54169 | 0.00054381 |  | 0.62817 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.53951
| 1000 | 1000 | 2024-11-19 11:57 | 0.63556 | 0 | 0.54004 | 0.0005481 | 1 | #N/A | 0.5395 | 0.5395 | 0.53949 | 0.53949 | 0.54224
| 1180000 | 1003 | 2024-11-19 11:58 | 0.62628 | 0 | 0.54171 | 0.00054453 |  | 0.6282 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.53953
| 1000 | 1000 | 2024-11-19 11:59 | 0.63555 | 0 | 0.54005 | 0.00054424 | 1 | #N/A | 0.53952 | 0.53951 | 0.5395 | 0.53949 | 0.54223
| 1194000 | 995 | 2024-11-19 11:59 | 0.62602 | 0 | 0.54171 | 0.00054269 |  | 0.62792 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.53954
| 1000 | 1000 | 2024-11-19 12:00 | 0.63557 | 0 | 0.54007 | 0.00054344 | 1 | #N/A | 0.53953 | 0.53953 | 0.53952 | 0.53952 | 0.54224
| 1200000 | 1000 | 2024-11-19 12:01 | 0.62584 | 0 | 0.5417 | 0.00055461 |  | 0.62781 | 0.54225 | 0.54225 | 0.54226 | 0.54226 | 0.53948
| 1000 | 1000 | 2024-11-19 12:02 | 0.63554 | 0 | 0.54 | 0.00055408 | 1 | #N/A | 0.53946 | 0.53945 | 0.53944 | 0.53943 | 0.54222
| 1206000 | 1005 | 2024-11-19 12:02 | 0.6257 | 0 | 0.54168 | 0.00055958 |  | 0.62767 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53944
| 1000 | 1000 | 2024-11-19 12:03 | 0.63554 | 0 | 0.53998 | 0.0005588 | 1 | #N/A | 0.53943 | 0.53942 | 0.53941 | 0.5394 | 0.54221
| 1212000 | 1010 | 2024-11-19 12:04 | 0.62558 | 0 | 0.54167 | 0.00055934 |  | 0.62754 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53944
| 1000 | 1000 | 2024-11-19 12:05 | 0.63553 | 0 | 0.53998 | 0.00055749 | 1 | #N/A | 0.53944 | 0.53943 | 0.53942 | 0.53941 | 0.54221
| 1240000 | 992 | 2024-11-19 12:05 | 0.62517 | 0 | 0.54167 | 0.00056465 |  | 0.62716 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.53942
| 1000 | 1000 | 2024-11-19 12:06 | 0.63555 | 0 | 0.53996 | 0.00056524 | 1 | #N/A | 0.53941 | 0.5394 | 0.53939 | 0.53939 | 0.54222
| 1245000 | 996 | 2024-11-19 12:07 | 0.62514 | 0 | 0.54169 | 0.00056471 |  | 0.62712 | 0.54223 | 0.54225 | 0.54226 | 0.54226 | 0.53943
| 1000 | 1000 | 2024-11-19 12:08 | 0.63557 | 0 | 0.53997 | 0.00056756 | 1 | #N/A | 0.53942 | 0.53941 | 0.5394 | 0.5394 | 0.54224
| 1250000 | 1000 | 2024-11-19 12:08 | 0.62504 | 0 | 0.54169 | 0.00057666 |  | 0.62707 | 0.54225 | 0.54227 | 0.54227 | 0.54228 | 0.53938
| 1000 | 1000 | 2024-11-19 12:09 | 0.63554 | 0 | 0.53993 | 0.00057179 | 1 | #N/A | 0.53937 | 0.53936 | 0.53935 | 0.53935 | 0.54222
| 1255000 | 1004 | 2024-11-19 12:10 | 0.62497 | 0 | 0.54166 | 0.00057407 |  | 0.62699 | 0.54223 | 0.54223 | 0.54224 | 0.54224 | 0.53937
| 1000 | 1000 | 2024-11-19 12:11 | 0.63554 | 0 | 0.53991 | 0.00057674 | 1 | #N/A | 0.53935 | 0.53934 | 0.53933 | 0.53933 | 0.54222
| 1275000 | 1003 | 2024-11-19 12:11 | 0.6248 | 0 | 0.54167 | 0.00058028 |  | 0.62684 | 0.54223 | 0.54224 | 0.54226 | 0.54226 | 0.53935
| 1000 | 1000 | 2024-11-19 12:12 | 0.63556 | 0 | 0.53991 | 0.00058239 | 1 | #N/A | 0.53934 | 0.53932 | 0.53932 | 0.53931 | 0.54224
| 1275000 | 1003 | 2024-11-19 12:13 | 0.6248 | 0 | 0.54168 | 0.00058166 |  | 0.62685 | 0.54225 | 0.54226 | 0.54227 | 0.54228 | 0.53936
| 1000 | 1000 | 2024-11-19 12:14 | 0.63555 | 0 | 0.53991 | 0.000578 | 1 | #N/A | 0.53935 | 0.53934 | 0.53933 | 0.53932 | 0.54222
| 1300000 | 1001 | 2024-11-19 12:14 | 0.62462 | 0 | 0.54165 | 0.00059011 |  | 0.62669 | 0.54223 | 0.54223 | 0.54224 | 0.54225 | 0.53929
| 1000 | 1000 | 2024-11-19 12:15 | 0.63555 | 0 | 0.53986 | 0.00059169 | 1 | #N/A | 0.53928 | 0.53927 | 0.53926 | 0.53925 | 0.54222
| 1300000 | 1001 | 2024-11-19 12:16 | 0.62462 | 0 | 0.54166 | 0.00059425 |  | 0.6267 | 0.54223 | 0.54224 | 0.54226 | 0.54228 | 0.53928
| 1000 | 1000 | 2024-11-19 12:16 | 0.63554 | 0 | 0.53986 | 0.00059016 | 1 | #N/A | 0.53928 | 0.53928 | 0.53927 | 0.53926 | 0.54222
| 1300000 | 1001 | 2024-11-19 12:17 | 0.62462 | 0 | 0.54165 | 0.00059117 |  | 0.6267 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53929
| 1000 | 1000 | 2024-11-19 12:18 | 0.63555 | 0 | 0.53986 | 0.00059233 | 1 | #N/A | 0.53928 | 0.53927 | 0.53926 | 0.53925 | 0.54223
| 1320000 | 1001 | 2024-11-19 12:19 | 0.62457 | 0 | 0.54164 | 0.00060087 |  | 0.62669 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53924
| 1000 | 1000 | 2024-11-19 12:19 | 0.63553 | 0 | 0.53981 | 0.00059967 | 1 | #N/A | 0.53922 | 0.53921 | 0.5392 | 0.5392 | 0.54221
| 1332000 | 999 | 2024-11-19 12:20 | 0.6246 | 0 | 0.54164 | 0.0005972 |  | 0.6267 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53925
| 1000 | 1000 | 2024-11-19 12:21 | 0.63556 | 0 | 0.53982 | 0.00060099 | 1 | #N/A | 0.53924 | 0.53923 | 0.53922 | 0.53921 | 0.54223
| 1336000 | 1002 | 2024-11-19 12:22 | 0.62462 | 0 | 0.54165 | 0.00060059 |  | 0.62673 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.53924
| 1000 | 1000 | 2024-11-19 12:22 | 0.63556 | 0 | 0.53983 | 0.00060279 | 1 | #N/A | 0.53923 | 0.53923 | 0.53922 | 0.53921 | 0.54224
| 1350000 | 999 | 2024-11-19 12:23 | 0.62464 | 0 | 0.54165 | 0.00061215 |  | 0.6268 | 0.54225 | 0.54226 | 0.54227 | 0.54227 | 0.5392
| 1000 | 1000 | 2024-11-19 12:24 | 0.63554 | 0 | 0.53979 | 0.00060726 | 1 | #N/A | 0.53919 | 0.53918 | 0.53918 | 0.53917 | 0.54222
| 1365000 | 1001 | 2024-11-19 12:25 | 0.62474 | 0 | 0.54163 | 0.0006117 |  | 0.62689 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53918
| 1000 | 1000 | 2024-11-19 12:25 | 0.63556 | 0 | 0.53977 | 0.0006145 | 1 | #N/A | 0.53917 | 0.53916 | 0.53915 | 0.53914 | 0.54223
| 1365000 | 1001 | 2024-11-19 12:26 | 0.62474 | 0 | 0.54164 | 0.00061337 |  | 0.6269 | 0.54224 | 0.54224 | 0.54226 | 0.54226 | 0.53918
| 1000 | 1000 | 2024-11-19 12:27 | 0.63557 | 0 | 0.53977 | 0.00061643 | 1 | #N/A | 0.53917 | 0.53916 | 0.53915 | 0.53914 | 0.54224
| 1375000 | 1001 | 2024-11-19 12:27 | 0.62485 | 0 | 0.54164 | 0.00061127 |  | 0.627 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.53919
| 1000 | 1000 | 2024-11-19 12:28 | 0.63557 | 0 | 0.53979 | 0.00061343 | 1 | #N/A | 0.53918 | 0.53917 | 0.53917 | 0.53916 | 0.54224
| 1400000 | 1001 | 2024-11-19 12:29 | 0.62511 | 0 | 0.54163 | 0.00062462 |  | 0.62729 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.53913
| 1000 | 1000 | 2024-11-19 12:30 | 0.63557 | 0 | 0.53975 | 0.00062546 | 1 | #N/A | 0.53912 | 0.53912 | 0.53912 | 0.53912 | 0.54225
| 1400000 | 1001 | 2024-11-19 12:30 | 0.62513 | 0 | 0.54164 | 0.00062152 |  | 0.62732 | 0.54226 | 0.54226 | 0.54227 | 0.54227 | 0.53916
| 1000 | 1000 | 2024-11-19 12:31 | 0.63557 | 0 | 0.53975 | 0.00062301 | 1 | #N/A | 0.53915 | 0.53913 | 0.53912 | 0.53912 | 0.54225
| 1416000 | 1003 | 2024-11-19 12:32 | 0.62531 | 0 | 0.54163 | 0.00063634 |  | 0.62754 | 0.54226 | 0.54227 | 0.54227 | 0.54228 | 0.53909
| 1000 | 1000 | 2024-11-19 12:33 | 0.63558 | 0 | 0.5397 | 0.00063625 | 1 | #N/A | 0.53908 | 0.53907 | 0.53906 | 0.53905 | 0.54225
| 1416000 | 1003 | 2024-11-19 12:33 | 0.62533 | 0 | 0.54164 | 0.00063252 |  | 0.62757 | 0.54226 | 0.54227 | 0.54228 | 0.54229 | 0.53911
| 1000 | 1000 | 2024-11-19 12:34 | 0.63556 | 0 | 0.53971 | 0.00063113 | 1 | #N/A | 0.5391 | 0.53909 | 0.53907 | 0.53906 | 0.54224
| 1430000 | 1001 | 2024-11-19 12:35 | 0.6255 | 0 | 0.54161 | 0.00064533 |  | 0.62777 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.53903
| 1000 | 1000 | 2024-11-19 12:36 | 0.63556 | 0 | 0.53965 | 0.00064549 | 1 | #N/A | 0.53902 | 0.53901 | 0.539 | 0.53899 | 0.54223
| 1440000 | 996 | 2024-11-19 12:36 | 0.62572 | 0 | 0.54161 | 0.00064088 |  | 0.62797 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.53904
| 1000 | 1000 | 2024-11-19 12:37 | 0.63556 | 0 | 0.53966 | 0.00064272 | 1 | #N/A | 0.53903 | 0.53902 | 0.53902 | 0.53901 | 0.54223
| 1440000 | 996 | 2024-11-19 12:38 | 0.62572 | 0 | 0.54161 | 0.00064102 |  | 0.62798 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.53905
| 1000 | 1000 | 2024-11-19 12:39 | 0.63556 | 0 | 0.53966 | 0.00064203 | 1 | #N/A | 0.53903 | 0.53902 | 0.53902 | 0.53901 | 0.54223
| 1450000 | 986 | 2024-11-19 12:39 | 0.6259 | 0 | 0.5416 | 0.00064881 |  | 0.62818 | 0.54224 | 0.54225 | 0.54225 | 0.54227 | 0.53901
| 1000 | 1000 | 2024-11-19 12:40 | 0.63555 | 0 | 0.53963 | 0.00064827 | 1 | #N/A | 0.539 | 0.53899 | 0.53898 | 0.53896 | 0.54222
| 1475000 | 1003 | 2024-11-19 12:41 | 0.6264 | 0 | 0.54158 | 0.00066491 |  | 0.62875 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53892
| 1000 | 1000 | 2024-11-19 12:42 | 0.63555 | 0 | 0.53956 | 0.00066393 | 1 | #N/A | 0.53891 | 0.5389 | 0.53889 | 0.53888 | 0.54222
| 1480000 | 999 | 2024-11-19 12:42 | 0.62656 | 0 | 0.54158 | 0.00065618 |  | 0.62887 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53895
| 1000 | 1000 | 2024-11-19 12:43 | 0.63555 | 0 | 0.53959 | 0.0006583 | 1 | #N/A | 0.53894 | 0.53894 | 0.53892 | 0.53892 | 0.54222
| 1494000 | 996 | 2024-11-19 12:44 | 0.62686 | 0 | 0.54158 | 0.00067375 |  | 0.62924 | 0.54223 | 0.54225 | 0.54225 | 0.54226 | 0.53888
| 1000 | 1000 | 2024-11-19 12:44 | 0.63554 | 0 | 0.53953 | 0.00067237 | 1 | #N/A | 0.53887 | 0.53887 | 0.53885 | 0.53884 | 0.54222
| 1500000 | 1000 | 2024-11-19 12:45 | 0.62707 | 0 | 0.54158 | 0.00066704 |  | 0.62942 | 0.54223 | 0.54225 | 0.54225 | 0.54226 | 0.53891
| 1000 | 1000 | 2024-11-19 12:46 | 0.63557 | 0 | 0.53956 | 0.00067011 | 1 | #N/A | 0.5389 | 0.53889 | 0.53889 | 0.53888 | 0.54224
| 1506000 | 1004 | 2024-11-19 12:47 | 0.62718 | 0 | 0.54158 | 0.00067945 |  | 0.62957 | 0.54225 | 0.54225 | 0.54226 | 0.54227 | 0.53886
| 1000 | 1000 | 2024-11-19 12:47 | 0.63555 | 0 | 0.53952 | 0.00067533 | 1 | #N/A | 0.53886 | 0.53885 | 0.53884 | 0.53883 | 0.54222
| 1515000 | 1010 | 2024-11-19 12:48 | 0.62746 | 0 | 0.54157 | 0.0006754 |  | 0.62984 | 0.54223 | 0.54225 | 0.54225 | 0.54226 | 0.53887
| 1000 | 1000 | 2024-11-19 12:49 | 0.63555 | 0 | 0.53952 | 0.00067597 | 1 | #N/A | 0.53886 | 0.53885 | 0.53883 | 0.53883 | 0.54222
| 1540000 | 1001 | 2024-11-19 12:50 | 0.62815 | 0 | 0.54155 | 0.00069126 |  | 0.6306 | 0.54223 | 0.54223 | 0.54225 | 0.54226 | 0.53878
| 1000 | 1000 | 2024-11-19 12:50 | 0.63555 | 0 | 0.53945 | 0.00069369 | 1 | #N/A | 0.53877 | 0.53875 | 0.53875 | 0.53874 | 0.54222
| 1540000 | 1001 | 2024-11-19 12:51 | 0.62815 | 0 | 0.54156 | 0.00069272 |  | 0.63059 | 0.54223 | 0.54224 | 0.54226 | 0.54226 | 0.53878
| 1000 | 1000 | 2024-11-19 12:52 | 0.63555 | 0 | 0.53946 | 0.00069182 | 1 | #N/A | 0.53877 | 0.53876 | 0.53876 | 0.53876 | 0.54222
| 1550000 | 992 | 2024-11-19 12:53 | 0.62848 | 0 | 0.54155 | 0.00069361 |  | 0.63093 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53878
| 1000 | 1000 | 2024-11-19 12:53 | 0.63554 | 0 | 0.53944 | 0.0006934 | 1 | #N/A | 0.53877 | 0.53875 | 0.53875 | 0.53874 | 0.54222
| 1560000 | 1001 | 2024-11-19 12:54 | 0.62879 | 0 | 0.54154 | 0.00069352 |  | 0.63124 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.53876
| 1000 | 1000 | 2024-11-19 12:55 | 0.63556 | 0 | 0.53944 | 0.00069794 | 1 | #N/A | 0.53875 | 0.53875 | 0.53873 | 0.53873 | 0.54223
| 1575000 | 1008 | 2024-11-19 12:56 | 0.62928 | 0 | 0.54154 | 0.00070735 |  | 0.63179 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.53872
| 1000 | 1000 | 2024-11-19 12:56 | 0.63554 | 0 | 0.53939 | 0.00070661 | 1 | #N/A | 0.5387 | 0.53869 | 0.53868 | 0.53867 | 0.54222
| 1575000 | 1008 | 2024-11-19 12:57 | 0.62928 | 0 | 0.54154 | 0.00070638 |  | 0.63178 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53871
| 1000 | 1000 | 2024-11-19 12:58 | 0.63555 | 0 | 0.53939 | 0.00070793 | 1 | #N/A | 0.5387 | 0.53869 | 0.53868 | 0.53867 | 0.54222
| 1592000 | 995 | 2024-11-19 12:58 | 0.62988 | 0 | 0.54153 | 0.00070817 |  | 0.63239 | 0.54224 | 0.54224 | 0.54224 | 0.54225 | 0.5387
| 1000 | 1000 | 2024-11-19 12:59 | 0.63554 | 0 | 0.53938 | 0.00070742 | 1 | #N/A | 0.53869 | 0.53868 | 0.53867 | 0.53866 | 0.54221
| 1600000 | 1000 | 2024-11-19 13:00 | 0.63019 | 0 | 0.54154 | 0.00070548 |  | 0.63268 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53872
| 1000 | 1000 | 2024-11-19 13:01 | 0.63555 | 0 | 0.5394 | 0.00070559 | 1 | #N/A | 0.53871 | 0.5387 | 0.53869 | 0.53869 | 0.54223
| 1608000 | 1005 | 2024-11-19 13:01 | 0.63043 | 0 | 0.54153 | 0.00071708 |  | 0.63297 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53866
| 1000 | 1000 | 2024-11-19 13:02 | 0.63555 | 0 | 0.53936 | 0.00071599 | 1 | #N/A | 0.53865 | 0.53864 | 0.53863 | 0.53863 | 0.54222
| 1620000 | 999 | 2024-11-19 13:03 | 0.63088 | 0 | 0.54152 | 0.00071631 |  | 0.63341 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53866
| 1000 | 1000 | 2024-11-19 13:04 | 0.63557 | 0 | 0.53936 | 0.00071936 | 1 | #N/A | 0.53865 | 0.53865 | 0.53864 | 0.53863 | 0.54224
| 1625000 | 1001 | 2024-11-19 13:04 | 0.63104 | 0 | 0.54153 | 0.00072403 |  | 0.63362 | 0.54224 | 0.54225 | 0.54225 | 0.54227 | 0.53863
| 1000 | 1000 | 2024-11-19 13:05 | 0.63556 | 0 | 0.53933 | 0.0007256 | 1 | #N/A | 0.53862 | 0.53861 | 0.5386 | 0.53859 | 0.54223
| 1650000 | 1001 | 2024-11-19 13:06 | 0.63199 | 0 | 0.54153 | 0.00072438 |  | 0.63457 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.53863
| 1000 | 1000 | 2024-11-19 13:07 | 0.63557 | 0 | 0.53933 | 0.00072738 | 1 | #N/A | 0.53862 | 0.53861 | 0.5386 | 0.5386 | 0.54224
| 1650000 | 1001 | 2024-11-19 13:07 | 0.63199 | 0 | 0.54153 | 0.00072756 |  | 0.63456 | 0.54225 | 0.54226 | 0.54226 | 0.54226 | 0.53862
| 1000 | 1000 | 2024-11-19 13:08 | 0.63557 | 0 | 0.53934 | 0.00072675 | 1 | #N/A | 0.53862 | 0.53862 | 0.53861 | 0.5386 | 0.54225
| 1660000 | 996 | 2024-11-19 13:09 | 0.63236 | 0 | 0.54153 | 0.00073089 |  | 0.63496 | 0.54225 | 0.54226 | 0.54227 | 0.54228 | 0.53861
| 1000 | 1000 | 2024-11-19 13:10 | 0.63556 | 0 | 0.53932 | 0.00072955 | 1 | #N/A | 0.5386 | 0.53859 | 0.53859 | 0.53858 | 0.54224
| 1670000 | 1002 | 2024-11-19 13:10 | 0.63268 | 0 | 0.54151 | 0.00074124 |  | 0.63532 | 0.54224 | 0.54225 | 0.54225 | 0.54227 | 0.53855
| 1000 | 1000 | 2024-11-19 13:11 | 0.63555 | 0 | 0.53926 | 0.00074116 | 1 | #N/A | 0.53854 | 0.53852 | 0.53851 | 0.53851 | 0.54223
| 1680000 | 994 | 2024-11-19 13:12 | 0.63303 | 0 | 0.54151 | 0.00074813 |  | 0.6357 | 0.54223 | 0.54226 | 0.54227 | 0.54228 | 0.53852
| 1000 | 1000 | 2024-11-19 13:13 | 0.63555 | 0 | 0.53924 | 0.00074669 | 1 | #N/A | 0.53851 | 0.5385 | 0.53849 | 0.53848 | 0.54223
| 1700000 | 1003 | 2024-11-19 13:13 | 0.63375 | 0 | 0.54149 | 0.00076013 |  | 0.63645 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.53845
| 1000 | 1000 | 2024-11-19 13:14 | 0.63555 | 0 | 0.53919 | 0.00075848 | 1 | #N/A | 0.53844 | 0.53843 | 0.53843 | 0.53842 | 0.54222
| 1700000 | 1003 | 2024-11-19 13:15 | 0.63374 | 0 | 0.54149 | 0.00075764 |  | 0.63645 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53846
| 1000 | 1000 | 2024-11-19 13:16 | 0.63555 | 0 | 0.53919 | 0.0007583 | 1 | #N/A | 0.53845 | 0.53844 | 0.53843 | 0.53842 | 0.54222
| 1716000 | 1001 | 2024-11-19 13:16 | 0.6343 | 0 | 0.54147 | 0.00076705 |  | 0.63702 | 0.54223 | 0.54223 | 0.54224 | 0.54225 | 0.5384
| 1000 | 1000 | 2024-11-19 13:17 | 0.63556 | 0 | 0.53915 | 0.00076927 | 1 | #N/A | 0.5384 | 0.53839 | 0.53838 | 0.53838 | 0.54223
| 1716000 | 1001 | 2024-11-19 13:18 | 0.6343 | 0 | 0.54149 | 0.00076935 |  | 0.63704 | 0.54224 | 0.54226 | 0.54226 | 0.54227 | 0.53841
| 1000 | 1000 | 2024-11-19 13:18 | 0.63556 | 0 | 0.53916 | 0.00076842 | 1 | #N/A | 0.5384 | 0.5384 | 0.53838 | 0.53838 | 0.54223
| 1728000 | 1008 | 2024-11-19 13:19 | 0.63473 | 0 | 0.54148 | 0.00077299 |  | 0.6375 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.53839
| 1000 | 1000 | 2024-11-19 13:20 | 0.63556 | 0 | 0.53913 | 0.0007747 | 1 | #N/A | 0.53837 | 0.53836 | 0.53835 | 0.53834 | 0.54223
| 1750000 | 1001 | 2024-11-19 13:21 | 0.63546 | 0 | 0.54147 | 0.00078769 |  | 0.63828 | 0.54225 | 0.54226 | 0.54226 | 0.54226 | 0.53832
| 1000 | 1000 | 2024-11-19 13:21 | 0.63555 | 0 | 0.53908 | 0.00078665 | 1 | #N/A | 0.53831 | 0.53829 | 0.53829 | 0.53828 | 0.54223
| 1750000 | 1001 | 2024-11-19 13:22 | 0.63545 | 0 | 0.54146 | 0.00078596 |  | 0.63826 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.53832
| 1000 | 1000 | 2024-11-19 13:23 | 0.63556 | 0 | 0.53908 | 0.0007875 | 1 | #N/A | 0.53831 | 0.5383 | 0.53829 | 0.53828 | 0.54223
| 1750000 | 1001 | 2024-11-19 13:24 | 0.63546 | 0 | 0.54147 | 0.00078441 |  | 0.63827 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.53833
| 1000 | 1000 | 2024-11-19 13:24 | 0.63556 | 0 | 0.53909 | 0.00078593 | 1 | #N/A | 0.53832 | 0.53831 | 0.5383 | 0.53829 | 0.54223
| 1770000 | 1003 | 2024-11-19 13:25 | 0.63611 | 0 | 0.54146 | 0.00079695 |  | 0.63896 | 0.54225 | 0.54226 | 0.54226 | 0.54227 | 0.53827
| 1000 | 1000 | 2024-11-19 13:26 | 0.63556 | 0 | 0.53905 | 0.00079662 | 1 | #N/A | 0.53827 | 0.53826 | 0.53825 | 0.53824 | 0.54224
| 1775000 | 994 | 2024-11-19 13:26 | 0.63624 | 0 | 0.54146 | 0.00080435 |  | 0.63912 | 0.54225 | 0.54226 | 0.54226 | 0.54227 | 0.53824
| 1000 | 1000 | 2024-11-19 13:27 | 0.63554 | 0 | 0.53902 | 0.00079945 | 1 | #N/A | 0.53823 | 0.53823 | 0.53822 | 0.5382 | 0.54221
| 1800000 | 999 | 2024-11-19 13:28 | 0.63702 | 0 | 0.54143 | 0.00081104 |  | 0.63992 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.53818
| 1000 | 1000 | 2024-11-19 13:29 | 0.63554 | 0 | 0.53897 | 0.00081136 | 1 | #N/A | 0.53817 | 0.53816 | 0.53815 | 0.53815 | 0.54222
| 1800000 | 999 | 2024-11-19 13:29 | 0.63701 | 0 | 0.54143 | 0.00081134 |  | 0.63991 | 0.54223 | 0.54223 | 0.54225 | 0.54226 | 0.53819
| 1000 | 1000 | 2024-11-19 13:30 | 0.63556 | 0 | 0.53898 | 0.00081389 | 1 | #N/A | 0.53817 | 0.53817 | 0.53816 | 0.53815 | 0.54223
| 1800000 | 999 | 2024-11-19 13:31 | 0.63703 | 0 | 0.54145 | 0.00081147 |  | 0.63993 | 0.54224 | 0.54225 | 0.54227 | 0.54228 | 0.53821
| 1000 | 1000 | 2024-11-19 13:32 | 0.63557 | 0 | 0.53899 | 0.00081114 | 1 | #N/A | 0.5382 | 0.53819 | 0.53818 | 0.53817 | 0.54224
| 1820000 | 1001 | 2024-11-19 13:32 | 0.63763 | 0 | 0.54144 | 0.00081463 |  | 0.64054 | 0.54225 | 0.54225 | 0.54226 | 0.54226 | 0.53818
| 1000 | 1000 | 2024-11-19 13:33 | 0.63559 | 0 | 0.53898 | 0.00081995 | 1 | #N/A | 0.53818 | 0.53816 | 0.53815 | 0.53814 | 0.54226
| 1820000 | 1001 | 2024-11-19 13:34 | 0.63763 | 0 | 0.54145 | 0.00081727 |  | 0.64056 | 0.54226 | 0.54227 | 0.54227 | 0.54228 | 0.53818
| 1000 | 1000 | 2024-11-19 13:35 | 0.63558 | 0 | 0.53897 | 0.00081864 | 1 | #N/A | 0.53817 | 0.53816 | 0.53815 | 0.53814 | 0.54225
| 1848000 | 1001 | 2024-11-19 13:35 | 0.6383 | 0 | 0.54143 | 0.0008364 |  | 0.64131 | 0.54226 | 0.54226 | 0.54226 | 0.54227 | 0.53808
| 1000 | 1000 | 2024-11-19 13:36 | 0.63556 | 0 | 0.53889 | 0.0008359 | 1 | #N/A | 0.53807 | 0.53806 | 0.53804 | 0.53803 | 0.54223
| 1850000 | 999 | 2024-11-19 13:37 | 0.63834 | 0 | 0.54141 | 0.00083789 |  | 0.64135 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.53806
| 1000 | 1000 | 2024-11-19 13:38 | 0.63556 | 0 | 0.53887 | 0.00083894 | 1 | #N/A | 0.53805 | 0.53804 | 0.53803 | 0.53802 | 0.54223
| 1860000 | 992 | 2024-11-19 13:38 | 0.63859 | 0 | 0.54142 | 0.00083939 |  | 0.64161 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.53806
| 1000 | 1000 | 2024-11-19 13:39 | 0.63555 | 0 | 0.53887 | 0.00083938 | 1 | #N/A | 0.53805 | 0.53803 | 0.53802 | 0.53801 | 0.54223
| 1875000 | 1000 | 2024-11-19 13:40 | 0.63888 | 0 | 0.54139 | 0.00085014 |  | 0.64193 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53799
| 1000 | 1000 | 2024-11-19 13:41 | 0.63555 | 0 | 0.53882 | 0.0008503 | 1 | #N/A | 0.53798 | 0.53797 | 0.53797 | 0.53796 | 0.54222
| 1875000 | 1000 | 2024-11-19 13:41 | 0.63886 | 0 | 0.5414 | 0.00085032 |  | 0.64192 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53799
| 1000 | 1000 | 2024-11-19 13:42 | 0.63554 | 0 | 0.53882 | 0.00085012 | 1 | #N/A | 0.53799 | 0.53797 | 0.53796 | 0.53795 | 0.54222
| 1890000 | 1008 | 2024-11-19 13:43 | 0.63916 | 0 | 0.54139 | 0.00085798 |  | 0.64224 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53795
| 1000 | 1000 | 2024-11-19 13:43 | 0.63554 | 0 | 0.53879 | 0.00085652 | 1 | #N/A | 0.53794 | 0.53794 | 0.53793 | 0.53792 | 0.54221
| 1900000 | 1007 | 2024-11-19 13:44 | 0.63931 | 0 | 0.54138 | 0.00086523 |  | 0.64242 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53792
| 1000 | 1000 | 2024-11-19 13:45 | 0.63555 | 0 | 0.53876 | 0.00086599 | 1 | #N/A | 0.53791 | 0.5379 | 0.53788 | 0.53787 | 0.54222
| 1900000 | 1007 | 2024-11-19 13:46 | 0.6393 | 0 | 0.54137 | 0.00086543 |  | 0.6424 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.53791
| 1000 | 1000 | 2024-11-19 13:46 | 0.63556 | 0 | 0.53876 | 0.00086776 | 1 | #N/A | 0.5379 | 0.53789 | 0.53788 | 0.53788 | 0.54223
| 1925000 | 1001 | 2024-11-19 13:47 | 0.63968 | 0 | 0.54137 | 0.00087611 |  | 0.64282 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.53787
| 1000 | 1000 | 2024-11-19 13:48 | 0.63557 | 0 | 0.53872 | 0.00087878 | 1 | #N/A | 0.53786 | 0.53785 | 0.53784 | 0.53783 | 0.54224
| 1925000 | 1001 | 2024-11-19 13:49 | 0.63969 | 0 | 0.54138 | 0.00087561 |  | 0.64284 | 0.54225 | 0.54225 | 0.54226 | 0.54227 | 0.53788
| 1000 | 1000 | 2024-11-19 13:49 | 0.63558 | 0 | 0.53873 | 0.00087859 | 1 | #N/A | 0.53787 | 0.53786 | 0.53785 | 0.53785 | 0.54225
| 1950000 | 1001 | 2024-11-19 13:50 | 0.63992 | 0 | 0.54137 | 0.00089362 |  | 0.64313 | 0.54226 | 0.54226 | 0.54226 | 0.54227 | 0.5378
| 1000 | 1000 | 2024-11-19 13:51 | 0.63557 | 0 | 0.53867 | 0.00089333 | 1 | #N/A | 0.53778 | 0.53778 | 0.53777 | 0.53776 | 0.54224
| 1950000 | 1001 | 2024-11-19 13:52 | 0.63992 | 0 | 0.54136 | 0.00089292 |  | 0.64313 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.53779
| 1000 | 1000 | 2024-11-19 13:52 | 0.63555 | 0 | 0.53866 | 0.0008918 | 1 | #N/A | 0.53778 | 0.53777 | 0.53776 | 0.53775 | 0.54222
| 1950000 | 1001 | 2024-11-19 13:53 | 0.63991 | 0 | 0.54135 | 0.00088955 |  | 0.64311 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.5378
| 1000 | 1000 | 2024-11-19 13:54 | 0.63555 | 0 | 0.53866 | 0.00089099 | 1 | #N/A | 0.53779 | 0.53777 | 0.53776 | 0.53775 | 0.54222
| 1976000 | 988 | 2024-11-19 13:55 | 0.6401 | 0 | 0.54135 | 0.00089935 |  | 0.64334 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53775
| 1000 | 1000 | 2024-11-19 13:55 | 0.63556 | 0 | 0.53862 | 0.00090184 | 1 | #N/A | 0.53774 | 0.53772 | 0.53772 | 0.53771 | 0.54223
| 1984000 | 992 | 2024-11-19 13:56 | 0.64015 | 0 | 0.54136 | 0.00090076 |  | 0.64339 | 0.54224 | 0.54225 | 0.54227 | 0.54227 | 0.53775
| 1000 | 1000 | 2024-11-19 13:57 | 0.63558 | 0 | 0.53863 | 0.00090412 | 1 | #N/A | 0.53774 | 0.53773 | 0.53772 | 0.53771 | 0.54225
| 1992000 | 996 | 2024-11-19 13:57 | 0.64012 | 0 | 0.54135 | 0.00091224 |  | 0.6434 | 0.54226 | 0.54226 | 0.54227 | 0.54227 | 0.5377
| 1000 | 1000 | 2024-11-19 13:58 | 0.63556 | 0 | 0.53859 | 0.00090983 | 1 | #N/A | 0.53769 | 0.53768 | 0.53768 | 0.53767 | 0.54223
| 2000000 | 1000 | 2024-11-19 13:59 | 0.64012 | 0 | 0.54134 | 0.00091515 |  | 0.64342 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.53768
| 1000 | 1000 | 2024-11-19 14:00 | 0.63553 | 0 | 0.53857 | 0.00091116 | 1 | #N/A | 0.53767 | 0.53766 | 0.53765 | 0.53764 | 0.54221
| 2008000 | 1004 | 2024-11-19 14:00 | 0.64011 | 0 | 0.54132 | 0.00091523 |  | 0.6434 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53766
| 1000 | 1000 | 2024-11-19 14:01 | 0.63556 | 0 | 0.53856 | 0.00091958 | 1 | #N/A | 0.53765 | 0.53764 | 0.53763 | 0.53763 | 0.54224
| 2025000 | 999 | 2024-11-19 14:02 | 0.64002 | 0 | 0.54132 | 0.00093382 |  | 0.64338 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.53758
| 1000 | 1000 | 2024-11-19 14:03 | 0.63555 | 0 | 0.53849 | 0.00093354 | 1 | #N/A | 0.53757 | 0.53756 | 0.53755 | 0.53755 | 0.54222
| 2025000 | 999 | 2024-11-19 14:03 | 0.64002 | 0 | 0.54131 | 0.00093068 |  | 0.64336 | 0.54223 | 0.54223 | 0.54224 | 0.54225 | 0.53758
| 1000 | 1000 | 2024-11-19 14:04 | 0.63555 | 0 | 0.53849 | 0.00093442 | 1 | #N/A | 0.53757 | 0.53756 | 0.53756 | 0.53755 | 0.54223
| 2040000 | 1003 | 2024-11-19 14:05 | 0.63993 | 0 | 0.54131 | 0.00094058 |  | 0.64331 | 0.54224 | 0.54225 | 0.54225 | 0.54227 | 0.53755
| 1000 | 1000 | 2024-11-19 14:06 | 0.63554 | 0 | 0.53846 | 0.00093816 | 1 | #N/A | 0.53754 | 0.53753 | 0.53752 | 0.53751 | 0.54222
| 2040000 | 1003 | 2024-11-19 14:06 | 0.63994 | 0 | 0.54131 | 0.00093658 |  | 0.64331 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53756
| 1000 | 1000 | 2024-11-19 14:07 | 0.63554 | 0 | 0.53847 | 0.00093749 | 1 | #N/A | 0.53755 | 0.53754 | 0.53753 | 0.53752 | 0.54222
| 2075000 | 996 | 2024-11-19 14:08 | 0.63955 | 0 | 0.54129 | 0.00096033 |  | 0.643 | 0.54223 | 0.54224 | 0.54226 | 0.54226 | 0.53745
| 1000 | 1000 | 2024-11-19 14:09 | 0.63554 | 0 | 0.53838 | 0.00095902 | 1 | #N/A | 0.53743 | 0.53743 | 0.53742 | 0.53742 | 0.54222
| 2075000 | 996 | 2024-11-19 14:09 | 0.63955 | 0 | 0.54129 | 0.00095766 |  | 0.64299 | 0.54223 | 0.54223 | 0.54225 | 0.54226 | 0.53746
| 1000 | 1000 | 2024-11-19 14:10 | 0.63555 | 0 | 0.53839 | 0.00095821 | 1 | #N/A | 0.53744 | 0.53743 | 0.53742 | 0.53741 | 0.54222
| 2075000 | 996 | 2024-11-19 14:11 | 0.63954 | 0 | 0.54128 | 0.00096058 |  | 0.64299 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53744
| 1000 | 1000 | 2024-11-19 14:12 | 0.63554 | 0 | 0.53838 | 0.00095935 | 1 | #N/A | 0.53743 | 0.53742 | 0.53741 | 0.5374 | 0.54221
| 2100000 | 1001 | 2024-11-19 14:12 | 0.6392 | 0 | 0.54128 | 0.00096495 |  | 0.64266 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53742
| 1000 | 1000 | 2024-11-19 14:13 | 0.63557 | 0 | 0.53836 | 0.00096968 | 1 | #N/A | 0.53741 | 0.5374 | 0.53739 | 0.53739 | 0.54224
| 2100000 | 1001 | 2024-11-19 14:14 | 0.63918 | 0 | 0.54129 | 0.00096969 |  | 0.64266 | 0.54225 | 0.54225 | 0.54226 | 0.54226 | 0.53741
| 1000 | 1000 | 2024-11-19 14:14 | 0.63556 | 0 | 0.53835 | 0.00097078 | 1 | #N/A | 0.5374 | 0.53738 | 0.53738 | 0.53737 | 0.54224
| 2100000 | 1001 | 2024-11-19 14:15 | 0.63918 | 0 | 0.54128 | 0.00096869 |  | 0.64266 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.53741
| 1000 | 1000 | 2024-11-19 14:16 | 0.63554 | 0 | 0.53835 | 0.00096811 | 1 | #N/A | 0.53739 | 0.53738 | 0.53737 | 0.53736 | 0.54222
| 2125000 | 1003 | 2024-11-19 14:17 | 0.63868 | 0 | 0.54126 | 0.00098191 |  | 0.64221 | 0.54223 | 0.54224 | 0.54224 | 0.54226 | 0.53733
| 1000 | 1000 | 2024-11-19 14:17 | 0.63554 | 0 | 0.53829 | 0.00098309 | 1 | #N/A | 0.53732 | 0.53731 | 0.5373 | 0.53729 | 0.54222
| 2125000 | 1003 | 2024-11-19 14:18 | 0.63865 | 0 | 0.54126 | 0.00098551 |  | 0.64218 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.53731
| 1000 | 1000 | 2024-11-19 14:19 | 0.63556 | 0 | 0.53829 | 0.00098569 | 1 | #N/A | 0.53731 | 0.53731 | 0.5373 | 0.53729 | 0.54223
| 2145000 | 1001 | 2024-11-19 14:20 | 0.63819 | 0 | 0.54126 | 0.0010016 |  | 0.64178 | 0.54224 | 0.54226 | 0.54227 | 0.54227 | 0.53725
| 1000 | 1000 | 2024-11-19 14:20 | 0.63556 | 0 | 0.53823 | 0.0010016 | 1 | #N/A | 0.53724 | 0.53723 | 0.53723 | 0.53722 | 0.54224
| 2145000 | 1001 | 2024-11-19 14:21 | 0.63817 | 0 | 0.54125 | 0.0010027 |  | 0.64176 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.53724
| 1000 | 1000 | 2024-11-19 14:22 | 0.63555 | 0 | 0.53822 | 0.0010008 | 1 | #N/A | 0.53723 | 0.53722 | 0.53722 | 0.53721 | 0.54222
| 2160000 | 1008 | 2024-11-19 14:23 | 0.63783 | 0 | 0.54124 | 0.0010024 |  | 0.64141 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53723
| 1000 | 1000 | 2024-11-19 14:23 | 0.63558 | 0 | 0.53822 | 0.0010076 | 1 | #N/A | 0.53722 | 0.53721 | 0.5372 | 0.5372 | 0.54225
| 2160000 | 1008 | 2024-11-19 14:24 | 0.63784 | 0 | 0.54126 | 0.0010053 |  | 0.64145 | 0.54226 | 0.54226 | 0.54227 | 0.54228 | 0.53724
| 1000 | 1000 | 2024-11-19 14:25 | 0.63556 | 0 | 0.53822 | 0.0010035 | 1 | #N/A | 0.53723 | 0.53722 | 0.53721 | 0.5372 | 0.54223
| 2184000 | 1001 | 2024-11-19 14:26 | 0.63713 | 0 | 0.54124 | 0.0010186 |  | 0.64077 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.53716
| 1000 | 1000 | 2024-11-19 14:26 | 0.63557 | 0 | 0.53816 | 0.0010211 | 1 | #N/A | 0.53715 | 0.53715 | 0.53713 | 0.53713 | 0.54224
| 2184000 | 1001 | 2024-11-19 14:27 | 0.63715 | 0 | 0.54125 | 0.0010205 |  | 0.64079 | 0.54226 | 0.54226 | 0.54227 | 0.54228 | 0.53716
| 1000 | 1000 | 2024-11-19 14:28 | 0.63558 | 0 | 0.53817 | 0.0010196 | 1 | #N/A | 0.53715 | 0.53715 | 0.53714 | 0.53714 | 0.54224
| 2200000 | 1001 | 2024-11-19 14:28 | 0.6367 | 0 | 0.54124 | 0.0010187 |  | 0.64033 | 0.54225 | 0.54225 | 0.54226 | 0.54227 | 0.53716
| 1000 | 1000 | 2024-11-19 14:29 | 0.6356 | 0 | 0.53817 | 0.0010216 | 1 | #N/A | 0.53716 | 0.53715 | 0.53715 | 0.53714 | 0.54226
| 2200000 | 1001 | 2024-11-19 14:30 | 0.63671 | 0 | 0.54125 | 0.0010211 |  | 0.64034 | 0.54226 | 0.54227 | 0.54228 | 0.54229 | 0.53717
| 1000 | 1000 | 2024-11-19 14:31 | 0.6356 | 0 | 0.53818 | 0.0010222 | 1 | #N/A | 0.53716 | 0.53716 | 0.53716 | 0.53715 | 0.54227
| 2220000 | 999 | 2024-11-19 14:31 | 0.63597 | 0 | 0.54124 | 0.0010486 |  | 0.63971 | 0.54228 | 0.54229 | 0.54229 | 0.54229 | 0.53704
| 1000 | 1000 | 2024-11-19 14:32 | 0.63558 | 0 | 0.53807 | 0.0010459 | 1 | #N/A | 0.53703 | 0.53703 | 0.53702 | 0.53701 | 0.54225
| 2220000 | 999 | 2024-11-19 14:33 | 0.63596 | 0 | 0.54122 | 0.0010463 |  | 0.6397 | 0.54226 | 0.54226 | 0.54227 | 0.54228 | 0.53704
| 1000 | 1000 | 2024-11-19 14:34 | 0.63559 | 0 | 0.53806 | 0.0010488 | 1 | #N/A | 0.53702 | 0.53702 | 0.537 | 0.537 | 0.54225
| 2240000 | 1008 | 2024-11-19 14:34 | 0.63527 | 0 | 0.54122 | 0.0010585 |  | 0.63906 | 0.54226 | 0.54227 | 0.54228 | 0.54229 | 0.53698
| 1000 | 1000 | 2024-11-19 14:35 | 0.63556 | 0 | 0.53801 | 0.0010561 | 1 | #N/A | 0.53697 | 0.53696 | 0.53695 | 0.53694 | 0.54223
| 2250000 | 999 | 2024-11-19 14:36 | 0.63486 | 0 | 0.54118 | 0.0010674 |  | 0.63867 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.53691
| 1000 | 1000 | 2024-11-19 14:37 | 0.63556 | 0 | 0.53796 | 0.0010687 | 1 | #N/A | 0.5369 | 0.53689 | 0.53688 | 0.53687 | 0.54223
| 2250000 | 999 | 2024-11-19 14:37 | 0.63483 | 0 | 0.54118 | 0.0010705 |  | 0.63865 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.5369
| 1000 | 1000 | 2024-11-19 14:38 | 0.63555 | 0 | 0.53795 | 0.0010691 | 1 | #N/A | 0.53689 | 0.53688 | 0.53688 | 0.53686 | 0.54222
| 2275000 | 1001 | 2024-11-19 14:39 | 0.6339 | 0 | 0.54117 | 0.0010857 |  | 0.63777 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.53682
| 1000 | 1000 | 2024-11-19 14:40 | 0.63556 | 0 | 0.53789 | 0.0010866 | 1 | #N/A | 0.53682 | 0.53681 | 0.53679 | 0.53678 | 0.54223
| 2275000 | 1001 | 2024-11-19 14:40 | 0.63389 | 0 | 0.54116 | 0.0010855 |  | 0.63776 | 0.54224 | 0.54224 | 0.54225 | 0.54225 | 0.53682
| 1000 | 1000 | 2024-11-19 14:41 | 0.63555 | 0 | 0.53788 | 0.0010864 | 1 | #N/A | 0.53681 | 0.5368 | 0.53679 | 0.53678 | 0.54223
| 2280000 | 1007 | 2024-11-19 14:42 | 0.63367 | 0 | 0.54115 | 0.0010921 |  | 0.63756 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53678
| 1000 | 1000 | 2024-11-19 14:43 | 0.63554 | 0 | 0.53785 | 0.0010905 | 1 | #N/A | 0.53677 | 0.53676 | 0.53675 | 0.53675 | 0.54221
| 2310000 | 1001 | 2024-11-19 14:43 | 0.63246 | 0 | 0.54112 | 0.0011135 |  | 0.63642 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53667
| 1000 | 1000 | 2024-11-19 14:44 | 0.63553 | 0 | 0.53775 | 0.0011137 | 1 | #N/A | 0.53665 | 0.53664 | 0.53663 | 0.53663 | 0.54221
| 2310000 | 1001 | 2024-11-19 14:45 | 0.63248 | 0 | 0.54112 | 0.0011087 |  | 0.63643 | 0.54222 | 0.54222 | 0.54223 | 0.54224 | 0.53668
| 1000 | 1000 | 2024-11-19 14:45 | 0.63553 | 0 | 0.53776 | 0.0011111 | 1 | #N/A | 0.53667 | 0.53666 | 0.53664 | 0.53663 | 0.54221
| 2325000 | 992 | 2024-11-19 14:46 | 0.63183 | 0 | 0.54111 | 0.0011236 |  | 0.63583 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.53661
| 1000 | 1000 | 2024-11-19 14:47 | 0.63552 | 0 | 0.53769 | 0.0011249 | 1 | #N/A | 0.53659 | 0.53658 | 0.53656 | 0.53655 | 0.54219
| 2325000 | 992 | 2024-11-19 14:48 | 0.63183 | 0 | 0.5411 | 0.0011246 |  | 0.63582 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.5366
| 1000 | 1000 | 2024-11-19 14:48 | 0.63554 | 0 | 0.5377 | 0.0011271 | 1 | #N/A | 0.53659 | 0.53658 | 0.53657 | 0.53656 | 0.54221
| 2340000 | 1014 | 2024-11-19 14:49 | 0.63125 | 0 | 0.5411 | 0.0011289 |  | 0.63525 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.53659
| 1000 | 1000 | 2024-11-19 14:50 | 0.63554 | 0 | 0.5377 | 0.0011283 | 1 | #N/A | 0.53658 | 0.53658 | 0.53657 | 0.53656 | 0.54221
| 2360000 | 1003 | 2024-11-19 14:51 | 0.63043 | 0 | 0.54109 | 0.0011412 |  | 0.63448 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.53652
| 1000 | 1000 | 2024-11-19 14:51 | 0.63554 | 0 | 0.53764 | 0.0011444 | 1 | #N/A | 0.5365 | 0.53649 | 0.53649 | 0.53648 | 0.54221
| 2360000 | 1003 | 2024-11-19 14:52 | 0.63045 | 0 | 0.54109 | 0.0011388 |  | 0.63449 | 0.54222 | 0.54223 | 0.54223 | 0.54225 | 0.53654
| 1000 | 1000 | 2024-11-19 14:53 | 0.63555 | 0 | 0.53765 | 0.0011416 | 1 | #N/A | 0.53653 | 0.53652 | 0.53651 | 0.5365 | 0.54222
| 2375000 | 1007 | 2024-11-19 14:54 | 0.62984 | 0 | 0.54109 | 0.0011471 |  | 0.6339 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.5365
| 1000 | 1000 | 2024-11-19 14:54 | 0.63554 | 0 | 0.53763 | 0.0011484 | 1 | #N/A | 0.53649 | 0.53648 | 0.53647 | 0.53647 | 0.54222
| 2375000 | 1007 | 2024-11-19 14:55 | 0.62985 | 0 | 0.54109 | 0.0011453 |  | 0.6339 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53651
| 1000 | 1000 | 2024-11-19 14:56 | 0.63557 | 0 | 0.53764 | 0.0011488 | 1 | #N/A | 0.5365 | 0.5365 | 0.53649 | 0.53648 | 0.54224
| 2388000 | 995 | 2024-11-19 14:56 | 0.62935 | 0 | 0.5411 | 0.0011527 |  | 0.63342 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.53649
| 1000 | 1000 | 2024-11-19 14:57 | 0.63557 | 0 | 0.53762 | 0.0011539 | 1 | #N/A | 0.53648 | 0.53647 | 0.53647 | 0.53646 | 0.54224
| 2400000 | 1000 | 2024-11-19 14:58 | 0.62884 | 0 | 0.5411 | 0.0011708 |  | 0.63298 | 0.54225 | 0.54227 | 0.54228 | 0.54229 | 0.53642
| 1000 | 1000 | 2024-11-19 14:59 | 0.63556 | 0 | 0.53756 | 0.0011678 | 1 | #N/A | 0.53641 | 0.5364 | 0.53639 | 0.53637 | 0.54223
| 2412000 | 1005 | 2024-11-19 14:59 | 0.62832 | 0 | 0.54107 | 0.0011821 |  | 0.63249 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.53635
| 1000 | 1000 | 2024-11-19 15:00 | 0.63556 | 0 | 0.53751 | 0.0011804 | 1 | #N/A | 0.53633 | 0.53633 | 0.53632 | 0.53632 | 0.54223
| 2412000 | 1005 | 2024-11-19 15:01 | 0.62832 | 0 | 0.54107 | 0.0011807 |  | 0.6325 | 0.54224 | 0.54225 | 0.54225 | 0.54225 | 0.53634
| 1000 | 1000 | 2024-11-19 15:02 | 0.63554 | 0 | 0.53749 | 0.0011816 | 1 | #N/A | 0.53633 | 0.53631 | 0.5363 | 0.53629 | 0.54222
| 2424000 | 1010 | 2024-11-19 15:02 | 0.62783 | 0 | 0.54104 | 0.0011885 |  | 0.63202 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.53629
| 1000 | 1000 | 2024-11-19 15:03 | 0.63553 | 0 | 0.53745 | 0.0011891 | 1 | #N/A | 0.53628 | 0.53627 | 0.53626 | 0.53625 | 0.54221
| 2436000 | 1015 | 2024-11-19 15:04 | 0.62736 | 0 | 0.54103 | 0.0011979 |  | 0.63159 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.53624
| 1000 | 1000 | 2024-11-19 15:05 | 0.63554 | 0 | 0.53741 | 0.0012001 | 1 | #N/A | 0.53623 | 0.53622 | 0.53621 | 0.5362 | 0.54222
| 2455000 | 982 | 2024-11-19 15:05 | 0.62665 | 0 | 0.54102 | 0.0012116 |  | 0.63091 | 0.54223 | 0.54223 | 0.54223 | 0.54224 | 0.53618
| 1000 | 1000 | 2024-11-19 15:06 | 0.63553 | 0 | 0.53736 | 0.001211 | 1 | #N/A | 0.53617 | 0.53616 | 0.53615 | 0.53614 | 0.54221
| 2465000 | 986 | 2024-11-19 15:07 | 0.62631 | 0 | 0.54102 | 0.0012161 |  | 0.6306 | 0.54222 | 0.54223 | 0.54224 | 0.54226 | 0.53616
| 1000 | 1000 | 2024-11-19 15:08 | 0.63553 | 0 | 0.53734 | 0.0012153 | 1 | #N/A | 0.53614 | 0.53613 | 0.53612 | 0.53611 | 0.5422
| 2475000 | 990 | 2024-11-19 15:08 | 0.62602 | 0 | 0.54102 | 0.0012078 |  | 0.63025 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.53619
| 1000 | 1000 | 2024-11-19 15:09 | 0.63557 | 0 | 0.53739 | 0.0012129 | 1 | #N/A | 0.53618 | 0.53618 | 0.53618 | 0.53618 | 0.54224
| 2485000 | 994 | 2024-11-19 15:10 | 0.62561 | 0 | 0.54103 | 0.0012347 |  | 0.62997 | 0.54226 | 0.54226 | 0.54228 | 0.54228 | 0.53609
| 1000 | 1000 | 2024-11-19 15:11 | 0.63554 | 0 | 0.53729 | 0.0012312 | 1 | #N/A | 0.53608 | 0.53606 | 0.53605 | 0.53604 | 0.54221
| 2490000 | 996 | 2024-11-19 15:11 | 0.62549 | 0 | 0.54102 | 0.001225 |  | 0.62978 | 0.54223 | 0.54225 | 0.54225 | 0.54225 | 0.53612
| 1000 | 1000 | 2024-11-19 15:12 | 0.63557 | 0 | 0.53733 | 0.001228 | 1 | #N/A | 0.53611 | 0.53611 | 0.5361 | 0.53609 | 0.54224
| 2500000 | 1000 | 2024-11-19 15:13 | 0.62512 | 0 | 0.54102 | 0.0012406 |  | 0.62949 | 0.54225 | 0.54226 | 0.54226 | 0.54227 | 0.53606
| 1000 | 1000 | 2024-11-19 15:14 | 0.63555 | 0 | 0.53727 | 0.0012375 | 1 | #N/A | 0.53605 | 0.53603 | 0.53602 | 0.53602 | 0.54222
| 2510000 | 1004 | 2024-11-19 15:14 | 0.6248 | 0 | 0.541 | 0.0012438 |  | 0.62916 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53602
| 1000 | 1000 | 2024-11-19 15:15 | 0.63554 | 0 | 0.53725 | 0.0012408 | 1 | #N/A | 0.53602 | 0.53601 | 0.53601 | 0.53599 | 0.54221
| 2515000 | 1006 | 2024-11-19 15:16 | 0.62466 | 0 | 0.541 | 0.0012452 |  | 0.62904 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53602
| 1000 | 1000 | 2024-11-19 15:16 | 0.63552 | 0 | 0.53723 | 0.0012433 | 1 | #N/A | 0.53601 | 0.53599 | 0.53598 | 0.53597 | 0.5422
| 2525000 | 1010 | 2024-11-19 15:17 | 0.62431 | 0 | 0.54097 | 0.0012569 |  | 0.62872 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.53594
| 1000 | 1000 | 2024-11-19 15:18 | 0.63552 | 0 | 0.53717 | 0.0012573 | 1 | #N/A | 0.53593 | 0.53592 | 0.53591 | 0.5359 | 0.5422
| 2550000 | 1003 | 2024-11-19 15:19 | 0.62366 | 0 | 0.54095 | 0.0012651 |  | 0.6281 | 0.54221 | 0.54221 | 0.54222 | 0.54223 | 0.53589
| 1000 | 1000 | 2024-11-19 15:19 | 0.63553 | 0 | 0.53714 | 0.0012669 | 1 | #N/A | 0.53588 | 0.53587 | 0.53586 | 0.53586 | 0.5422
| 2550000 | 1003 | 2024-11-19 15:20 | 0.62367 | 0 | 0.54096 | 0.0012646 |  | 0.62811 | 0.5422 | 0.54222 | 0.54223 | 0.54224 | 0.5359
| 1000 | 1000 | 2024-11-19 15:21 | 0.63552 | 0 | 0.53713 | 0.0012666 | 1 | #N/A | 0.53588 | 0.53587 | 0.53586 | 0.53585 | 0.5422
| 2550000 | 1003 | 2024-11-19 15:22 | 0.62367 | 0 | 0.54096 | 0.0012619 |  | 0.6281 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.53591
| 1000 | 1000 | 2024-11-19 15:22 | 0.63554 | 0 | 0.53715 | 0.0012663 | 1 | #N/A | 0.5359 | 0.53589 | 0.53588 | 0.53587 | 0.54222
| 2580000 | 989 | 2024-11-19 15:23 | 0.62303 | 0 | 0.54096 | 0.0012773 |  | 0.62749 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53586
| 1000 | 1000 | 2024-11-19 15:24 | 0.63556 | 0 | 0.53712 | 0.0012772 | 1 | #N/A | 0.53585 | 0.53585 | 0.53584 | 0.53583 | 0.54223
| 2580000 | 989 | 2024-11-19 15:25 | 0.62303 | 0 | 0.54097 | 0.0012811 |  | 0.62751 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.53585
| 1000 | 1000 | 2024-11-19 15:25 | 0.63557 | 0 | 0.53711 | 0.0012827 | 1 | #N/A | 0.53584 | 0.53583 | 0.53582 | 0.53582 | 0.54224
| 2600000 | 1001 | 2024-11-19 15:26 | 0.62261 | 0 | 0.54096 | 0.0012973 |  | 0.62715 | 0.54225 | 0.54226 | 0.54226 | 0.54227 | 0.53577
| 1000 | 1000 | 2024-11-19 15:27 | 0.63556 | 0 | 0.53705 | 0.0012972 | 1 | #N/A | 0.53576 | 0.53575 | 0.53574 | 0.53573 | 0.54223
| 2600000 | 1001 | 2024-11-19 15:27 | 0.62262 | 0 | 0.54096 | 0.0012982 |  | 0.62717 | 0.54224 | 0.54226 | 0.54227 | 0.54228 | 0.53577
| 1000 | 1000 | 2024-11-19 15:28 | 0.63554 | 0 | 0.53704 | 0.0012946 | 1 | #N/A | 0.53576 | 0.53575 | 0.53574 | 0.53573 | 0.54221
| 2600000 | 1001 | 2024-11-19 15:29 | 0.62259 | 0 | 0.54094 | 0.0012967 |  | 0.62713 | 0.54222 | 0.54224 | 0.54224 | 0.54226 | 0.53576
| 1000 | 1000 | 2024-11-19 15:30 | 0.63555 | 0 | 0.53703 | 0.0012986 | 1 | #N/A | 0.53575 | 0.53574 | 0.53573 | 0.53573 | 0.54223
| 2625000 | 1001 | 2024-11-19 15:30 | 0.62219 | 0 | 0.54093 | 0.0013121 |  | 0.62679 | 0.54224 | 0.54224 | 0.54225 | 0.54225 | 0.53568
| 1000 | 1000 | 2024-11-19 15:31 | 0.63553 | 0 | 0.53696 | 0.0013115 | 1 | #N/A | 0.53567 | 0.53565 | 0.53564 | 0.53563 | 0.54221
| 2625000 | 1001 | 2024-11-19 15:32 | 0.6222 | 0 | 0.54092 | 0.0013103 |  | 0.62678 | 0.54221 | 0.54223 | 0.54224 | 0.54225 | 0.53568
| 1000 | 1000 | 2024-11-19 15:33 | 0.63553 | 0 | 0.53697 | 0.00131 | 1 | #N/A | 0.53567 | 0.53567 | 0.53566 | 0.53564 | 0.54221
| 2640000 | 990 | 2024-11-19 15:33 | 0.62203 | 0 | 0.54091 | 0.0013153 |  | 0.62663 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.53565
| 1000 | 1000 | 2024-11-19 15:34 | 0.63554 | 0 | 0.53694 | 0.0013175 | 1 | #N/A | 0.53564 | 0.53563 | 0.53562 | 0.53561 | 0.54221
| 2650000 | 1007 | 2024-11-19 15:35 | 0.62202 | 0 | 0.54093 | 0.001309 |  | 0.62659 | 0.54223 | 0.54223 | 0.54224 | 0.54225 | 0.53569
| 1000 | 1000 | 2024-11-19 15:36 | 0.63556 | 0 | 0.53698 | 0.0013119 | 1 | #N/A | 0.53569 | 0.53568 | 0.53567 | 0.53566 | 0.54223
| 2664000 | 999 | 2024-11-19 15:36 | 0.62188 | 0 | 0.54093 | 0.0013248 |  | 0.62652 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.53563
| 1000 | 1000 | 2024-11-19 15:37 | 0.63554 | 0 | 0.53692 | 0.0013235 | 1 | #N/A | 0.53562 | 0.53561 | 0.53559 | 0.53558 | 0.54222
| 2672000 | 1002 | 2024-11-19 15:38 | 0.62181 | 0 | 0.54091 | 0.0013279 |  | 0.62646 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.5356
| 1000 | 1000 | 2024-11-19 15:39 | 0.63554 | 0 | 0.5369 | 0.0013286 | 1 | #N/A | 0.53559 | 0.53558 | 0.53557 | 0.53556 | 0.54222
| 2680000 | 1005 | 2024-11-19 15:39 | 0.62179 | 0 | 0.54091 | 0.0013301 |  | 0.62643 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53559
| 1000 | 1000 | 2024-11-19 15:40 | 0.63556 | 0 | 0.5369 | 0.0013342 | 1 | #N/A | 0.53558 | 0.53557 | 0.53556 | 0.53555 | 0.54224
| 2688000 | 1008 | 2024-11-19 15:41 | 0.62181 | 0 | 0.54092 | 0.001332 |  | 0.62647 | 0.54225 | 0.54225 | 0.54226 | 0.54227 | 0.5356
| 1000 | 1000 | 2024-11-19 15:42 | 0.63554 | 0 | 0.5369 | 0.0013311 | 1 | #N/A | 0.53559 | 0.53557 | 0.53556 | 0.53555 | 0.54222
| 2700000 | 999 | 2024-11-19 15:42 | 0.62174 | 0 | 0.5409 | 0.0013442 |  | 0.62643 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53552
| 1000 | 1000 | 2024-11-19 15:43 | 0.63553 | 0 | 0.53684 | 0.0013422 | 1 | #N/A | 0.53552 | 0.53551 | 0.5355 | 0.53549 | 0.54221
| 2700000 | 999 | 2024-11-19 15:44 | 0.62175 | 0 | 0.5409 | 0.0013445 |  | 0.62645 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53552
| 1000 | 1000 | 2024-11-19 15:44 | 0.63553 | 0 | 0.53684 | 0.001344 | 1 | #N/A | 0.53551 | 0.5355 | 0.53549 | 0.53548 | 0.54221
| 2730000 | 1001 | 2024-11-19 15:45 | 0.62187 | 0 | 0.54089 | 0.0013543 |  | 0.62661 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.53547
| 1000 | 1000 | 2024-11-19 15:46 | 0.63554 | 0 | 0.5368 | 0.0013547 | 1 | #N/A | 0.53546 | 0.53544 | 0.53543 | 0.53543 | 0.54221
| 2730000 | 1001 | 2024-11-19 15:47 | 0.62186 | 0 | 0.54088 | 0.0013528 |  | 0.62659 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53547
| 1000 | 1000 | 2024-11-19 15:47 | 0.63553 | 0 | 0.5368 | 0.0013534 | 1 | #N/A | 0.53546 | 0.53545 | 0.53544 | 0.53543 | 0.54221
| 2730000 | 1001 | 2024-11-19 15:48 | 0.62187 | 0 | 0.54088 | 0.001352 |  | 0.62659 | 0.54222 | 0.54223 | 0.54225 | 0.54225 | 0.53548
| 1000 | 1000 | 2024-11-19 15:49 | 0.63555 | 0 | 0.53681 | 0.0013545 | 1 | #N/A | 0.53547 | 0.53545 | 0.53545 | 0.53545 | 0.54222
| 2750000 | 1001 | 2024-11-19 15:50 | 0.62204 | 0 | 0.54089 | 0.001355 |  | 0.62676 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.53547
| 1000 | 1000 | 2024-11-19 15:50 | 0.63557 | 0 | 0.53682 | 0.0013582 | 1 | #N/A | 0.53547 | 0.53546 | 0.53546 | 0.53545 | 0.54225
| 2750000 | 1001 | 2024-11-19 15:51 | 0.62207 | 0 | 0.54091 | 0.0013603 |  | 0.62682 | 0.54226 | 0.54227 | 0.54228 | 0.54228 | 0.53547
| 1000 | 1000 | 2024-11-19 15:52 | 0.63556 | 0 | 0.53681 | 0.0013549 | 1 | #N/A | 0.53546 | 0.53546 | 0.53546 | 0.53545 | 0.54223
| 2775000 | 999 | 2024-11-19 15:53 | 0.62227 | 0 | 0.54088 | 0.0013759 |  | 0.62709 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.53537
| 1000 | 1000 | 2024-11-19 15:53 | 0.63553 | 0 | 0.53672 | 0.001373 | 1 | #N/A | 0.53536 | 0.53535 | 0.53534 | 0.53533 | 0.54221
| 2775000 | 999 | 2024-11-19 15:54 | 0.62229 | 0 | 0.54087 | 0.0013749 |  | 0.62711 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53537
| 1000 | 1000 | 2024-11-19 15:55 | 0.63553 | 0 | 0.53671 | 0.0013725 | 1 | #N/A | 0.53536 | 0.53535 | 0.53533 | 0.53532 | 0.5422
| 2800000 | 994 | 2024-11-19 15:56 | 0.62274 | 0 | 0.54085 | 0.0013731 |  | 0.62755 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.53536
| 1000 | 1000 | 2024-11-19 15:56 | 0.63553 | 0 | 0.5367 | 0.0013763 | 1 | #N/A | 0.53535 | 0.53533 | 0.53532 | 0.53531 | 0.54221
| 2800000 | 994 | 2024-11-19 15:57 | 0.62274 | 0 | 0.54087 | 0.0013735 |  | 0.62755 | 0.54222 | 0.54223 | 0.54225 | 0.54225 | 0.53537
| 1000 | 1000 | 2024-11-19 15:58 | 0.63556 | 0 | 0.53673 | 0.0013775 | 1 | #N/A | 0.53536 | 0.53535 | 0.53535 | 0.53534 | 0.54224
| 2800000 | 994 | 2024-11-19 15:58 | 0.62276 | 0 | 0.54088 | 0.0013762 |  | 0.62758 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.53537
| 1000 | 1000 | 2024-11-19 15:59 | 0.63556 | 0 | 0.53673 | 0.0013767 | 1 | #N/A | 0.53536 | 0.53535 | 0.53534 | 0.53534 | 0.54223
| 2820000 | 987 | 2024-11-19 16:00 | 0.6231 | 0 | 0.54086 | 0.0013893 |  | 0.62797 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.5353
| 1000 | 1000 | 2024-11-19 16:01 | 0.63554 | 0 | 0.53666 | 0.0013886 | 1 | #N/A | 0.53528 | 0.53528 | 0.53527 | 0.53525 | 0.54221
| 2840000 | 994 | 2024-11-19 16:01 | 0.62352 | 0 | 0.54084 | 0.0013988 |  | 0.62843 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53524
| 1000 | 1000 | 2024-11-19 16:02 | 0.63553 | 0 | 0.53662 | 0.0013979 | 1 | #N/A | 0.53523 | 0.53523 | 0.53521 | 0.53521 | 0.54221
| 2840000 | 994 | 2024-11-19 16:03 | 0.62355 | 0 | 0.54084 | 0.0013968 |  | 0.62845 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.53526
| 1000 | 1000 | 2024-11-19 16:04 | 0.63555 | 0 | 0.53663 | 0.0013985 | 1 | #N/A | 0.53524 | 0.53523 | 0.53522 | 0.53522 | 0.54222
| 2850000 | 1007 | 2024-11-19 16:04 | 0.62377 | 0 | 0.54084 | 0.0014043 |  | 0.62869 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53522
| 1000 | 1000 | 2024-11-19 16:05 | 0.63554 | 0 | 0.5366 | 0.0014036 | 1 | #N/A | 0.53521 | 0.5352 | 0.5352 | 0.53519 | 0.54222
| 2860000 | 1001 | 2024-11-19 16:06 | 0.62403 | 0 | 0.54084 | 0.0014102 |  | 0.62899 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.5352
| 1000 | 1000 | 2024-11-19 16:07 | 0.63552 | 0 | 0.53657 | 0.0014072 | 1 | #N/A | 0.53518 | 0.53517 | 0.53516 | 0.53515 | 0.5422
| 2860000 | 1001 | 2024-11-19 16:07 | 0.62401 | 0 | 0.54082 | 0.0014062 |  | 0.62895 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.53519
| 1000 | 1000 | 2024-11-19 16:08 | 0.63552 | 0 | 0.53657 | 0.0014076 | 1 | #N/A | 0.53518 | 0.53517 | 0.53516 | 0.53515 | 0.5422
| 2880000 | 1008 | 2024-11-19 16:09 | 0.62461 | 0 | 0.54082 | 0.0014046 |  | 0.62954 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.5352
| 1000 | 1000 | 2024-11-19 16:10 | 0.63554 | 0 | 0.53658 | 0.0014093 | 1 | #N/A | 0.53519 | 0.53518 | 0.53517 | 0.53516 | 0.54222
| 2880000 | 1008 | 2024-11-19 16:10 | 0.62464 | 0 | 0.54084 | 0.0014055 |  | 0.62959 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53522
| 1000 | 1000 | 2024-11-19 16:11 | 0.63555 | 0 | 0.5366 | 0.0014077 | 1 | #N/A | 0.53521 | 0.53519 | 0.53518 | 0.53517 | 0.54223
| 2900000 | 986 | 2024-11-19 16:12 | 0.6252 | 0 | 0.54083 | 0.0014207 |  | 0.63021 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.53515
| 1000 | 1000 | 2024-11-19 16:13 | 0.63552 | 0 | 0.53653 | 0.0014181 | 1 | #N/A | 0.53513 | 0.53512 | 0.5351 | 0.53509 | 0.5422
| 2900000 | 986 | 2024-11-19 16:13 | 0.62518 | 0 | 0.54081 | 0.00142 |  | 0.63017 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.53513
| 1000 | 1000 | 2024-11-19 16:14 | 0.63553 | 0 | 0.53652 | 0.0014211 | 1 | #N/A | 0.53512 | 0.53511 | 0.5351 | 0.53509 | 0.54221
| 2925000 | 1014 | 2024-11-19 16:15 | 0.62599 | 0 | 0.5408 | 0.0014294 |  | 0.63102 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53509
| 1000 | 1000 | 2024-11-19 16:15 | 0.63554 | 0 | 0.53649 | 0.0014307 | 1 | #N/A | 0.53507 | 0.53506 | 0.53506 | 0.53505 | 0.54222
| 2925000 | 1014 | 2024-11-19 16:16 | 0.62602 | 0 | 0.54081 | 0.0014259 |  | 0.63104 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53511
| 1000 | 1000 | 2024-11-19 16:17 | 0.63555 | 0 | 0.53651 | 0.0014285 | 1 | #N/A | 0.53509 | 0.53509 | 0.53508 | 0.53507 | 0.54222
| 2950000 | 1003 | 2024-11-19 16:18 | 0.62683 | 0 | 0.5408 | 0.0014463 |  | 0.63194 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53501
| 1000 | 1000 | 2024-11-19 16:18 | 0.63552 | 0 | 0.53642 | 0.0014448 | 1 | #N/A | 0.535 | 0.53498 | 0.53497 | 0.53496 | 0.5422
| 2950000 | 1003 | 2024-11-19 16:19 | 0.62682 | 0 | 0.54079 | 0.0014469 |  | 0.63191 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.535
| 1000 | 1000 | 2024-11-19 16:20 | 0.63553 | 0 | 0.53643 | 0.0014453 | 1 | #N/A | 0.53499 | 0.53499 | 0.53498 | 0.53497 | 0.54221
| 2950000 | 1003 | 2024-11-19 16:21 | 0.62683 | 0 | 0.54079 | 0.0014461 |  | 0.63193 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53501
| 1000 | 1000 | 2024-11-19 16:21 | 0.63554 | 0 | 0.53643 | 0.0014478 | 1 | #N/A | 0.53499 | 0.53498 | 0.53497 | 0.53497 | 0.54222
| 2973000 | 991 | 2024-11-19 16:22 | 0.62764 | 0 | 0.54078 | 0.0014572 |  | 0.63279 | 0.54223 | 0.54223 | 0.54224 | 0.54225 | 0.53495
| 1000 | 1000 | 2024-11-19 16:23 | 0.63552 | 0 | 0.53638 | 0.0014553 | 1 | #N/A | 0.53494 | 0.53493 | 0.53492 | 0.5349 | 0.5422
| 2982000 | 994 | 2024-11-19 16:24 | 0.62798 | 0 | 0.54077 | 0.0014574 |  | 0.63313 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.53494
| 1000 | 1000 | 2024-11-19 16:24 | 0.63552 | 0 | 0.53636 | 0.0014599 | 1 | #N/A | 0.53492 | 0.53491 | 0.5349 | 0.53488 | 0.5422
| 2991000 | 997 | 2024-11-19 16:25 | 0.62829 | 0 | 0.54076 | 0.0014627 |  | 0.63346 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.53491
| 1000 | 1000 | 2024-11-19 16:26 | 0.63554 | 0 | 0.53636 | 0.0014637 | 1 | #N/A | 0.5349 | 0.5349 | 0.53489 | 0.53488 | 0.54221
| 3000000 | 1000 | 2024-11-19 16:27 | 0.62866 | 0 | 0.54077 | 0.0014612 |  | 0.63383 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53493
| 1000 | 1000 | 2024-11-19 16:27 | 0.63555 | 0 | 0.53637 | 0.0014637 | 1 | #N/A | 0.53491 | 0.53491 | 0.5349 | 0.5349 | 0.54222
| 3009000 | 1003 | 2024-11-19 16:28 | 0.62895 | 0 | 0.54077 | 0.0014757 |  | 0.63418 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53486
| 1000 | 1000 | 2024-11-19 16:29 | 0.63554 | 0 | 0.53631 | 0.0014764 | 1 | #N/A | 0.53485 | 0.53484 | 0.53483 | 0.53481 | 0.54221
| 3018000 | 1006 | 2024-11-19 16:30 | 0.62927 | 0 | 0.54076 | 0.0014836 |  | 0.63452 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53482
| 1000 | 1000 | 2024-11-19 16:30 | 0.63553 | 0 | 0.53628 | 0.0014813 | 1 | #N/A | 0.53482 | 0.53481 | 0.5348 | 0.53479 | 0.54221
| 3027000 | 1009 | 2024-11-19 16:31 | 0.62964 | 0 | 0.54076 | 0.001479 |  | 0.63487 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53485
| 1000 | 1000 | 2024-11-19 16:32 | 0.63556 | 0 | 0.53631 | 0.0014792 | 1 | #N/A | 0.53484 | 0.53484 | 0.53483 | 0.53482 | 0.54223
| 3040000 | 988 | 2024-11-19 16:32 | 0.63014 | 0 | 0.54077 | 0.0014863 |  | 0.63541 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.53482
| 1000 | 1000 | 2024-11-19 16:33 | 0.63556 | 0 | 0.53629 | 0.0014859 | 1 | #N/A | 0.53481 | 0.53481 | 0.5348 | 0.53479 | 0.54223
| 3040000 | 988 | 2024-11-19 16:34 | 0.63014 | 0 | 0.54076 | 0.0014849 |  | 0.63539 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.53482
| 1000 | 1000 | 2024-11-19 16:35 | 0.63559 | 0 | 0.5363 | 0.0014898 | 1 | #N/A | 0.53481 | 0.53481 | 0.53481 | 0.5348 | 0.54226
| 3054000 | 1018 | 2024-11-19 16:35 | 0.6306 | 0 | 0.54076 | 0.0015054 |  | 0.63595 | 0.54226 | 0.54227 | 0.54226 | 0.54226 | 0.53474
| 1000 | 1000 | 2024-11-19 16:36 | 0.63553 | 0 | 0.53621 | 0.0015003 | 1 | #N/A | 0.53472 | 0.53471 | 0.5347 | 0.53469 | 0.54221
| 3080000 | 1001 | 2024-11-19 16:37 | 0.63153 | 0 | 0.54072 | 0.0015123 |  | 0.63692 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53467
| 1000 | 1000 | 2024-11-19 16:38 | 0.63553 | 0 | 0.53615 | 0.0015145 | 1 | #N/A | 0.53466 | 0.53464 | 0.53463 | 0.53462 | 0.54221
| 3080000 | 1001 | 2024-11-19 16:38 | 0.63152 | 0 | 0.54072 | 0.0015181 |  | 0.63692 | 0.54222 | 0.54224 | 0.54224 | 0.54225 | 0.53465
| 1000 | 1000 | 2024-11-19 16:39 | 0.63553 | 0 | 0.53614 | 0.0015196 | 1 | #N/A | 0.53463 | 0.53462 | 0.53461 | 0.53459 | 0.54221
| 3080000 | 1001 | 2024-11-19 16:40 | 0.63152 | 0 | 0.54073 | 0.0015187 |  | 0.63692 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53465
| 1000 | 1000 | 2024-11-19 16:41 | 0.63555 | 0 | 0.53614 | 0.0015202 | 1 | #N/A | 0.53464 | 0.53463 | 0.53462 | 0.5346 | 0.54222
| 3100000 | 992 | 2024-11-19 16:41 | 0.63222 | 0 | 0.54071 | 0.0015277 |  | 0.63766 | 0.54223 | 0.54224 | 0.54224 | 0.54224 | 0.5346
| 1000 | 1000 | 2024-11-19 16:42 | 0.63553 | 0 | 0.53609 | 0.0015273 | 1 | #N/A | 0.53458 | 0.53457 | 0.53456 | 0.53455 | 0.5422
| 3100000 | 992 | 2024-11-19 16:43 | 0.63223 | 0 | 0.54071 | 0.001528 |  | 0.63766 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.5346
| 1000 | 1000 | 2024-11-19 16:44 | 0.63553 | 0 | 0.5361 | 0.0015271 | 1 | #N/A | 0.53459 | 0.53458 | 0.53456 | 0.53455 | 0.54221
| 3125000 | 1000 | 2024-11-19 16:44 | 0.63304 | 0 | 0.54069 | 0.001543 |  | 0.63853 | 0.54221 | 0.54223 | 0.54224 | 0.54224 | 0.53451
| 1000 | 1000 | 2024-11-19 16:45 | 0.63554 | 0 | 0.53603 | 0.0015454 | 1 | #N/A | 0.5345 | 0.53449 | 0.53448 | 0.53447 | 0.54221
| 3125000 | 1000 | 2024-11-19 16:46 | 0.63302 | 0 | 0.54068 | 0.001546 |  | 0.63852 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.5345
| 1000 | 1000 | 2024-11-19 16:46 | 0.63554 | 0 | 0.53603 | 0.0015476 | 1 | #N/A | 0.53449 | 0.53448 | 0.53448 | 0.53447 | 0.54222
| 3150000 | 1008 | 2024-11-19 16:47 | 0.63378 | 0 | 0.54067 | 0.0015654 |  | 0.63938 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53441
| 1000 | 1000 | 2024-11-19 16:48 | 0.63551 | 0 | 0.53594 | 0.0015628 | 1 | #N/A | 0.53439 | 0.53438 | 0.53437 | 0.53436 | 0.54219
| 3150000 | 1008 | 2024-11-19 16:49 | 0.6338 | 0 | 0.54066 | 0.0015614 |  | 0.63938 | 0.5422 | 0.54222 | 0.54223 | 0.54224 | 0.53442
| 1000 | 1000 | 2024-11-19 16:49 | 0.63554 | 0 | 0.53595 | 0.0015671 | 1 | #N/A | 0.5344 | 0.53438 | 0.53437 | 0.53436 | 0.54221
| 3150000 | 1008 | 2024-11-19 16:50 | 0.63379 | 0 | 0.54068 | 0.0015642 |  | 0.63937 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.53442
| 1000 | 1000 | 2024-11-19 16:51 | 0.63554 | 0 | 0.53596 | 0.0015652 | 1 | #N/A | 0.53441 | 0.5344 | 0.53439 | 0.53438 | 0.54222
| 3180000 | 1007 | 2024-11-19 16:52 | 0.63466 | 0 | 0.54066 | 0.0015857 |  | 0.64033 | 0.54223 | 0.54225 | 0.54226 | 0.54226 | 0.53432
| 1000 | 1000 | 2024-11-19 16:52 | 0.63552 | 0 | 0.53588 | 0.0015802 | 1 | #N/A | 0.53431 | 0.5343 | 0.53429 | 0.53429 | 0.5422
| 3180000 | 1007 | 2024-11-19 16:53 | 0.63463 | 0 | 0.54065 | 0.0015837 |  | 0.64028 | 0.54221 | 0.54223 | 0.54223 | 0.54224 | 0.53431
| 1000 | 1000 | 2024-11-19 16:54 | 0.63554 | 0 | 0.53587 | 0.001585 | 1 | #N/A | 0.5343 | 0.53429 | 0.53429 | 0.53428 | 0.54221
| 3180000 | 1007 | 2024-11-19 16:55 | 0.63465 | 0 | 0.54066 | 0.0015829 |  | 0.64031 | 0.54222 | 0.54223 | 0.54225 | 0.54226 | 0.53433
| 1000 | 1000 | 2024-11-19 16:55 | 0.63553 | 0 | 0.53588 | 0.0015825 | 1 | #N/A | 0.53431 | 0.53431 | 0.5343 | 0.53428 | 0.54221
| 3200000 | 992 | 2024-11-19 16:56 | 0.63523 | 0 | 0.54065 | 0.0015878 |  | 0.64089 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.5343
| 1000 | 1000 | 2024-11-19 16:57 | 0.63557 | 0 | 0.53588 | 0.0015898 | 1 | #N/A | 0.53429 | 0.53429 | 0.53429 | 0.53428 | 0.54224
| 3200000 | 992 | 2024-11-19 16:58 | 0.63524 | 0 | 0.54067 | 0.0015874 |  | 0.64091 | 0.54225 | 0.54225 | 0.54226 | 0.54227 | 0.53432
| 1000 | 1000 | 2024-11-19 16:58 | 0.63556 | 0 | 0.53588 | 0.0015882 | 1 | #N/A | 0.53431 | 0.5343 | 0.53429 | 0.53428 | 0.54224
| 3225000 | 989 | 2024-11-19 16:59 | 0.63572 | 0 | 0.54064 | 0.0016155 |  | 0.64151 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.53417
| 1000 | 1000 | 2024-11-19 17:00 | 0.63555 | 0 | 0.53576 | 0.0016156 | 1 | #N/A | 0.53416 | 0.53415 | 0.53414 | 0.53412 | 0.54222
| 3240000 | 999 | 2024-11-19 17:00 | 0.63604 | 0 | 0.54062 | 0.0016233 |  | 0.64184 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53413
| 1000 | 1000 | 2024-11-19 17:01 | 0.63555 | 0 | 0.53574 | 0.0016219 | 1 | #N/A | 0.53413 | 0.53412 | 0.53411 | 0.5341 | 0.54222
| 3240000 | 999 | 2024-11-19 17:02 | 0.63602 | 0 | 0.54062 | 0.0016238 |  | 0.64182 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53412
| 1000 | 1000 | 2024-11-19 17:03 | 0.63555 | 0 | 0.53573 | 0.0016227 | 1 | #N/A | 0.53411 | 0.53411 | 0.53411 | 0.53409 | 0.54222
| 3250000 | 1001 | 2024-11-19 17:03 | 0.63615 | 0 | 0.54061 | 0.0016365 |  | 0.64201 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53406
| 1000 | 1000 | 2024-11-19 17:04 | 0.63555 | 0 | 0.53567 | 0.0016378 | 1 | #N/A | 0.53404 | 0.53403 | 0.53403 | 0.53403 | 0.54222
| 3250000 | 1001 | 2024-11-19 17:05 | 0.63617 | 0 | 0.54061 | 0.0016365 |  | 0.64204 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53406
| 1000 | 1000 | 2024-11-19 17:06 | 0.63552 | 0 | 0.53566 | 0.0016343 | 1 | #N/A | 0.53405 | 0.53404 | 0.53402 | 0.534 | 0.5422
| 3280000 | 984 | 2024-11-19 17:06 | 0.63662 | 0 | 0.54057 | 0.0016508 |  | 0.64251 | 0.54221 | 0.54222 | 0.54222 | 0.54223 | 0.53397
| 1000 | 1000 | 2024-11-19 17:07 | 0.63557 | 0 | 0.53561 | 0.0016571 | 1 | #N/A | 0.53396 | 0.53396 | 0.53395 | 0.53395 | 0.54224
| 3280000 | 984 | 2024-11-19 17:08 | 0.63663 | 0 | 0.5406 | 0.0016557 |  | 0.64256 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.53398
| 1000 | 1000 | 2024-11-19 17:09 | 0.63555 | 0 | 0.53561 | 0.0016534 | 1 | #N/A | 0.53397 | 0.53396 | 0.53395 | 0.53394 | 0.54222
| 3300000 | 1001 | 2024-11-19 17:09 | 0.63682 | 0 | 0.54058 | 0.0016694 |  | 0.64279 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.53391
| 1000 | 1000 | 2024-11-19 17:10 | 0.63558 | 0 | 0.53556 | 0.0016726 | 1 | #N/A | 0.5339 | 0.53389 | 0.53388 | 0.53387 | 0.54225
| 3300000 | 1001 | 2024-11-19 17:11 | 0.63681 | 0 | 0.5406 | 0.0016725 |  | 0.64281 | 0.54226 | 0.54226 | 0.54227 | 0.54228 | 0.53391
| 1000 | 1000 | 2024-11-19 17:12 | 0.63557 | 0 | 0.53555 | 0.0016714 | 1 | #N/A | 0.53389 | 0.53388 | 0.53387 | 0.53387 | 0.54224
| 3310000 | 993 | 2024-11-19 17:12 | 0.63684 | 0 | 0.54057 | 0.0016851 |  | 0.64288 | 0.54225 | 0.54225 | 0.54226 | 0.54227 | 0.53383
| 1000 | 1000 | 2024-11-19 17:13 | 0.63555 | 0 | 0.53549 | 0.0016841 | 1 | #N/A | 0.53382 | 0.53381 | 0.5338 | 0.53379 | 0.54222
| 3320000 | 996 | 2024-11-19 17:14 | 0.63692 | 0 | 0.54056 | 0.0016869 |  | 0.64296 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53381
| 1000 | 1000 | 2024-11-19 17:15 | 0.63557 | 0 | 0.53547 | 0.0016911 | 1 | #N/A | 0.5338 | 0.53378 | 0.53378 | 0.53377 | 0.54224
| 3330000 | 999 | 2024-11-19 17:15 | 0.63688 | 0 | 0.54055 | 0.0017115 |  | 0.64303 | 0.54225 | 0.54226 | 0.54227 | 0.54228 | 0.53371
| 1000 | 1000 | 2024-11-19 17:16 | 0.63553 | 0 | 0.53538 | 0.0017047 | 1 | #N/A | 0.5337 | 0.53368 | 0.53368 | 0.53367 | 0.5422
| 3340000 | 1002 | 2024-11-19 17:17 | 0.63691 | 0 | 0.54052 | 0.0017092 |  | 0.64304 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.53368
| 1000 | 1000 | 2024-11-19 17:17 | 0.63554 | 0 | 0.53536 | 0.0017115 | 1 | #N/A | 0.53366 | 0.53366 | 0.53365 | 0.53364 | 0.54221
| 3350000 | 1005 | 2024-11-19 17:18 | 0.63692 | 0 | 0.54051 | 0.0017166 |  | 0.64306 | 0.54221 | 0.54223 | 0.54223 | 0.54224 | 0.53365
| 1000 | 1000 | 2024-11-19 17:19 | 0.63556 | 0 | 0.53535 | 0.0017209 | 1 | #N/A | 0.53364 | 0.53363 | 0.53362 | 0.53362 | 0.54223
| 3360000 | 1008 | 2024-11-19 17:20 | 0.63691 | 0 | 0.54053 | 0.0017291 |  | 0.64311 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.53361
| 1000 | 1000 | 2024-11-19 17:20 | 0.63557 | 0 | 0.53532 | 0.0017311 | 1 | #N/A | 0.5336 | 0.53359 | 0.53358 | 0.53357 | 0.54224
| 3375000 | 999 | 2024-11-19 17:21 | 0.63678 | 0 | 0.54051 | 0.0017517 |  | 0.64307 | 0.54225 | 0.54226 | 0.54226 | 0.54228 | 0.5335
| 1000 | 1000 | 2024-11-19 17:22 | 0.63554 | 0 | 0.53522 | 0.0017484 | 1 | #N/A | 0.53349 | 0.53348 | 0.53347 | 0.53346 | 0.54222
| 3375000 | 999 | 2024-11-19 17:23 | 0.63677 | 0 | 0.54049 | 0.0017502 |  | 0.64305 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.53349
| 1000 | 1000 | 2024-11-19 17:23 | 0.63554 | 0 | 0.53521 | 0.0017502 | 1 | #N/A | 0.53348 | 0.53347 | 0.53345 | 0.53344 | 0.54221
| 3400000 | 1003 | 2024-11-19 17:24 | 0.63658 | 0 | 0.54047 | 0.0017637 |  | 0.6429 | 0.54222 | 0.54222 | 0.54223 | 0.54224 | 0.53341
| 1000 | 1000 | 2024-11-19 17:25 | 0.63556 | 0 | 0.53516 | 0.0017678 | 1 | #N/A | 0.5334 | 0.53339 | 0.53338 | 0.53338 | 0.54223
| 3400000 | 1003 | 2024-11-19 17:26 | 0.6366 | 0 | 0.54048 | 0.0017683 |  | 0.64295 | 0.54223 | 0.54224 | 0.54226 | 0.54226 | 0.53341
| 1000 | 1000 | 2024-11-19 17:26 | 0.63555 | 0 | 0.53515 | 0.0017688 | 1 | #N/A | 0.53339 | 0.53338 | 0.53337 | 0.53336 | 0.54222
| 3408000 | 994 | 2024-11-19 17:27 | 0.63649 | 0 | 0.54047 | 0.0017748 |  | 0.64285 | 0.54224 | 0.54224 | 0.54225 | 0.54225 | 0.53337
| 1000 | 1000 | 2024-11-19 17:28 | 0.63554 | 0 | 0.53512 | 0.0017738 | 1 | #N/A | 0.53336 | 0.53335 | 0.53334 | 0.53334 | 0.54221
| 3432000 | 1001 | 2024-11-19 17:29 | 0.63614 | 0 | 0.54044 | 0.0017939 |  | 0.64256 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53326
| 1000 | 1000 | 2024-11-19 17:29 | 0.63556 | 0 | 0.53504 | 0.0017976 | 1 | #N/A | 0.53325 | 0.53324 | 0.53324 | 0.53323 | 0.54223
| 3432000 | 1001 | 2024-11-19 17:30 | 0.63617 | 0 | 0.54045 | 0.001797 |  | 0.6426 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.53326
| 1000 | 1000 | 2024-11-19 17:31 | 0.63556 | 0 | 0.53504 | 0.0017966 | 1 | #N/A | 0.53326 | 0.53325 | 0.53325 | 0.53324 | 0.54223
| 3432000 | 1001 | 2024-11-19 17:31 | 0.63615 | 0 | 0.54045 | 0.0017972 |  | 0.64259 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.53326
| 1000 | 1000 | 2024-11-19 17:32 | 0.63556 | 0 | 0.53504 | 0.001798 | 1 | #N/A | 0.53325 | 0.53324 | 0.53323 | 0.53322 | 0.54223
| 3456000 | 1008 | 2024-11-19 17:33 | 0.6357 | 0 | 0.54043 | 0.0018191 |  | 0.64222 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.53316
| 1000 | 1000 | 2024-11-19 17:34 | 0.63556 | 0 | 0.53495 | 0.0018207 | 1 | #N/A | 0.53314 | 0.53314 | 0.53312 | 0.53311 | 0.54223
| 3456000 | 1008 | 2024-11-19 17:34 | 0.6357 | 0 | 0.54043 | 0.0018181 |  | 0.64221 | 0.54224 | 0.54224 | 0.54225 | 0.54225 | 0.53316
| 1000 | 1000 | 2024-11-19 17:35 | 0.63554 | 0 | 0.53495 | 0.0018161 | 1 | #N/A | 0.53315 | 0.53314 | 0.53313 | 0.53312 | 0.54221
| 3456000 | 1008 | 2024-11-19 17:36 | 0.63571 | 0 | 0.54042 | 0.0018139 |  | 0.64222 | 0.54223 | 0.54223 | 0.54224 | 0.54225 | 0.53317
| 1000 | 1000 | 2024-11-19 17:37 | 0.63553 | 0 | 0.53495 | 0.0018146 | 1 | #N/A | 0.53316 | 0.53314 | 0.53313 | 0.53311 | 0.54221
| 3480000 | 986 | 2024-11-19 17:37 | 0.63517 | 0 | 0.5404 | 0.0018351 |  | 0.64174 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.53305
| 1000 | 1000 | 2024-11-19 17:38 | 0.63555 | 0 | 0.53487 | 0.0018392 | 1 | #N/A | 0.53304 | 0.53303 | 0.53302 | 0.53302 | 0.54222
| 3500000 | 1001 | 2024-11-19 17:39 | 0.63463 | 0 | 0.54038 | 0.0018597 |  | 0.64128 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53295
| 1000 | 1000 | 2024-11-19 17:40 | 0.63554 | 0 | 0.53478 | 0.0018609 | 1 | #N/A | 0.53294 | 0.53292 | 0.53291 | 0.5329 | 0.54222
| 3500000 | 1001 | 2024-11-19 17:40 | 0.63462 | 0 | 0.54038 | 0.0018605 |  | 0.64128 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53294
| 1000 | 1000 | 2024-11-19 17:41 | 0.63554 | 0 | 0.53477 | 0.0018609 | 1 | #N/A | 0.53293 | 0.53291 | 0.5329 | 0.53289 | 0.54221
| 3500000 | 1001 | 2024-11-19 17:42 | 0.6346 | 0 | 0.54037 | 0.0018612 |  | 0.64125 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.53293
| 1000 | 1000 | 2024-11-19 17:43 | 0.63555 | 0 | 0.53477 | 0.0018619 | 1 | #N/A | 0.53292 | 0.53291 | 0.53291 | 0.5329 | 0.54222
| 3525000 | 987 | 2024-11-19 17:43 | 0.63381 | 0 | 0.54035 | 0.0018962 |  | 0.64059 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.53277
| 1000 | 1000 | 2024-11-19 17:44 | 0.63553 | 0 | 0.53464 | 0.0018917 | 1 | #N/A | 0.53276 | 0.53275 | 0.53275 | 0.53274 | 0.54221
| 3540000 | 1003 | 2024-11-19 17:45 | 0.63336 | 0 | 0.54033 | 0.0019008 |  | 0.64014 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53273
| 1000 | 1000 | 2024-11-19 17:46 | 0.63554 | 0 | 0.53461 | 0.0019013 | 1 | #N/A | 0.53272 | 0.53271 | 0.53271 | 0.5327 | 0.54221
| 3540000 | 1003 | 2024-11-19 17:46 | 0.63339 | 0 | 0.54034 | 0.0018994 |  | 0.64016 | 0.54223 | 0.54223 | 0.54224 | 0.54225 | 0.53274
| 1000 | 1000 | 2024-11-19 17:47 | 0.63556 | 0 | 0.53462 | 0.0019029 | 1 | #N/A | 0.53273 | 0.53272 | 0.53271 | 0.5327 | 0.54223
| 3550000 | 994 | 2024-11-19 17:48 | 0.63301 | 0 | 0.54034 | 0.0019156 |  | 0.63985 | 0.54224 | 0.54225 | 0.54225 | 0.54227 | 0.53267
| 1000 | 1000 | 2024-11-19 17:49 | 0.63554 | 0 | 0.53456 | 0.0019132 | 1 | #N/A | 0.53266 | 0.53265 | 0.53264 | 0.53263 | 0.54221
| 3550000 | 994 | 2024-11-19 17:49 | 0.63302 | 0 | 0.54032 | 0.0019134 |  | 0.63984 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53267
| 1000 | 1000 | 2024-11-19 17:50 | 0.63553 | 0 | 0.53455 | 0.0019138 | 1 | #N/A | 0.53266 | 0.53265 | 0.53264 | 0.53262 | 0.54221
| 3575000 | 1001 | 2024-11-19 17:51 | 0.63213 | 0 | 0.5403 | 0.0019331 |  | 0.63901 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53257
| 1000 | 1000 | 2024-11-19 17:51 | 0.63555 | 0 | 0.53448 | 0.0019345 | 1 | #N/A | 0.53257 | 0.53256 | 0.53254 | 0.53253 | 0.54222
| 3575000 | 1001 | 2024-11-19 17:52 | 0.63215 | 0 | 0.54031 | 0.0019342 |  | 0.63905 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.53257
| 1000 | 1000 | 2024-11-19 17:53 | 0.63553 | 0 | 0.53447 | 0.001934 | 1 | #N/A | 0.53256 | 0.53254 | 0.53253 | 0.53253 | 0.54221
| 3600000 | 1002 | 2024-11-19 17:54 | 0.63117 | 0 | 0.54028 | 0.0019592 |  | 0.63814 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53244
| 1000 | 1000 | 2024-11-19 17:54 | 0.63556 | 0 | 0.53438 | 0.0019649 | 1 | #N/A | 0.53243 | 0.53241 | 0.53241 | 0.5324 | 0.54224
| 3600000 | 1002 | 2024-11-19 17:55 | 0.63116 | 0 | 0.54029 | 0.0019598 |  | 0.63814 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.53245
| 1000 | 1000 | 2024-11-19 17:56 | 0.63556 | 0 | 0.53438 | 0.0019616 | 1 | #N/A | 0.53244 | 0.53242 | 0.53241 | 0.5324 | 0.54223
| 3600000 | 1002 | 2024-11-19 17:57 | 0.63115 | 0 | 0.54028 | 0.0019611 |  | 0.63812 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.53244
| 1000 | 1000 | 2024-11-19 17:57 | 0.63554 | 0 | 0.53438 | 0.0019593 | 1 | #N/A | 0.53243 | 0.53242 | 0.53242 | 0.53241 | 0.54222
| 3625000 | 986 | 2024-11-19 17:58 | 0.63013 | 0 | 0.54025 | 0.0019883 |  | 0.63719 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.5323
| 1000 | 1000 | 2024-11-19 17:59 | 0.63553 | 0 | 0.53426 | 0.0019858 | 1 | #N/A | 0.53229 | 0.53227 | 0.53227 | 0.53226 | 0.5422
| 3640000 | 1001 | 2024-11-19 18:00 | 0.62956 | 0 | 0.54024 | 0.001993 |  | 0.63663 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53227
| 1000 | 1000 | 2024-11-19 18:00 | 0.63556 | 0 | 0.53424 | 0.0019956 | 1 | #N/A | 0.53226 | 0.53225 | 0.53224 | 0.53224 | 0.54223
| 3640000 | 1001 | 2024-11-19 18:01 | 0.62958 | 0 | 0.54025 | 0.0019947 |  | 0.63665 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.53227
| 1000 | 1000 | 2024-11-19 18:02 | 0.63556 | 0 | 0.53424 | 0.0019966 | 1 | #N/A | 0.53226 | 0.53225 | 0.53224 | 0.53223 | 0.54223
| 3640000 | 1001 | 2024-11-19 18:02 | 0.62959 | 0 | 0.54026 | 0.001992 |  | 0.63666 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.53229
| 1000 | 1000 | 2024-11-19 18:03 | 0.63557 | 0 | 0.53426 | 0.0019963 | 1 | #N/A | 0.53228 | 0.53227 | 0.53226 | 0.53225 | 0.54224
| 3640000 | 1001 | 2024-11-19 18:04 | 0.62955 | 0 | 0.54027 | 0.001997 |  | 0.63664 | 0.54225 | 0.54226 | 0.54227 | 0.54227 | 0.53228
| 1000 | 1000 | 2024-11-19 18:05 | 0.63555 | 0 | 0.53425 | 0.001994 | 1 | #N/A | 0.53227 | 0.53226 | 0.53225 | 0.53225 | 0.54223
| 3680000 | 1012 | 2024-11-19 18:05 | 0.6278 | 0 | 0.54021 | 0.0020382 |  | 0.63503 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.53206
| 1000 | 1000 | 2024-11-19 18:06 | 0.63553 | 0 | 0.53406 | 0.0020379 | 1 | #N/A | 0.53205 | 0.53203 | 0.53202 | 0.53201 | 0.54222
| 3680000 | 1012 | 2024-11-19 18:07 | 0.62782 | 0 | 0.5402 | 0.0020365 |  | 0.63502 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53205
| 1000 | 1000 | 2024-11-19 18:08 | 0.63555 | 0 | 0.53407 | 0.0020377 | 1 | #N/A | 0.53204 | 0.53204 | 0.53203 | 0.53202 | 0.54222
| 3700000 | 999 | 2024-11-19 18:08 | 0.6269 | 0 | 0.54018 | 0.0020634 |  | 0.6342 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.53192
| 1000 | 1000 | 2024-11-19 18:09 | 0.63554 | 0 | 0.53396 | 0.0020655 | 1 | #N/A | 0.53191 | 0.5319 | 0.53189 | 0.53188 | 0.54222
| 3700000 | 999 | 2024-11-19 18:10 | 0.62688 | 0 | 0.54017 | 0.0020654 |  | 0.63418 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.53191
| 1000 | 1000 | 2024-11-19 18:11 | 0.63554 | 0 | 0.53395 | 0.0020645 | 1 | #N/A | 0.5319 | 0.53189 | 0.53189 | 0.53188 | 0.54221
| 3700000 | 999 | 2024-11-19 18:11 | 0.6269 | 0 | 0.54017 | 0.0020619 |  | 0.6342 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53192
| 1000 | 1000 | 2024-11-19 18:12 | 0.63552 | 0 | 0.53396 | 0.0020614 | 1 | #N/A | 0.53191 | 0.5319 | 0.53189 | 0.53188 | 0.5422
| 3720000 | 992 | 2024-11-19 18:13 | 0.62605 | 0 | 0.54015 | 0.0020783 |  | 0.6334 | 0.54221 | 0.54222 | 0.54224 | 0.54225 | 0.53184
| 1000 | 1000 | 2024-11-19 18:14 | 0.63553 | 0 | 0.53389 | 0.0020796 | 1 | #N/A | 0.53183 | 0.53181 | 0.5318 | 0.53179 | 0.54221
| 3735000 | 996 | 2024-11-19 18:14 | 0.62536 | 0 | 0.54014 | 0.0021028 |  | 0.63279 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53173
| 1000 | 1000 | 2024-11-19 18:15 | 0.63553 | 0 | 0.5338 | 0.0021007 | 1 | #N/A | 0.53171 | 0.5317 | 0.53169 | 0.53169 | 0.5422
| 3735000 | 996 | 2024-11-19 18:16 | 0.62532 | 0 | 0.54012 | 0.0021018 |  | 0.63275 | 0.54221 | 0.54222 | 0.54223 | 0.54223 | 0.53172
| 1000 | 1000 | 2024-11-19 18:17 | 0.6355 | 0 | 0.53379 | 0.0020987 | 1 | #N/A | 0.53171 | 0.53169 | 0.53168 | 0.53167 | 0.54218
| 3750000 | 1000 | 2024-11-19 18:17 | 0.62475 | 0 | 0.54011 | 0.0021044 |  | 0.63215 | 0.54219 | 0.54221 | 0.54222 | 0.54223 | 0.53169
| 1000 | 1000 | 2024-11-19 18:18 | 0.63555 | 0 | 0.53378 | 0.0021111 | 1 | #N/A | 0.53168 | 0.53167 | 0.53166 | 0.53165 | 0.54222
| 3765000 | 1004 | 2024-11-19 18:19 | 0.6241 | 0 | 0.54011 | 0.0021324 |  | 0.63161 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53158
| 1000 | 1000 | 2024-11-19 18:19 | 0.63553 | 0 | 0.53369 | 0.0021312 | 1 | #N/A | 0.53157 | 0.53156 | 0.53155 | 0.53154 | 0.54221
| 3765000 | 1004 | 2024-11-19 18:20 | 0.6241 | 0 | 0.5401 | 0.0021307 |  | 0.63162 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.53158
| 1000 | 1000 | 2024-11-19 18:21 | 0.63553 | 0 | 0.53368 | 0.0021306 | 1 | #N/A | 0.53157 | 0.53155 | 0.53154 | 0.53153 | 0.5422
| 3780000 | 1008 | 2024-11-19 18:22 | 0.6235 | 0 | 0.54009 | 0.0021358 |  | 0.63101 | 0.54221 | 0.54222 | 0.54223 | 0.54225 | 0.53155
| 1000 | 1000 | 2024-11-19 18:22 | 0.63555 | 0 | 0.53366 | 0.0021402 | 1 | #N/A | 0.53153 | 0.53152 | 0.53151 | 0.53151 | 0.54222
| 3800000 | 1007 | 2024-11-19 18:23 | 0.62271 | 0 | 0.54008 | 0.0021611 |  | 0.63031 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53144
| 1000 | 1000 | 2024-11-19 18:24 | 0.63554 | 0 | 0.53357 | 0.0021618 | 1 | #N/A | 0.53143 | 0.53141 | 0.5314 | 0.53139 | 0.54222
| 3800000 | 1007 | 2024-11-19 18:25 | 0.62268 | 0 | 0.54009 | 0.0021651 |  | 0.63028 | 0.54223 | 0.54225 | 0.54226 | 0.54226 | 0.53142
| 1000 | 1000 | 2024-11-19 18:25 | 0.63554 | 0 | 0.53357 | 0.0021636 | 1 | #N/A | 0.53142 | 0.53141 | 0.53139 | 0.53139 | 0.54222
| 3800000 | 1007 | 2024-11-19 18:26 | 0.62268 | 0 | 0.54008 | 0.0021639 |  | 0.63029 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.53142
| 1000 | 1000 | 2024-11-19 18:27 | 0.63553 | 0 | 0.53356 | 0.0021606 | 1 | #N/A | 0.53141 | 0.5314 | 0.5314 | 0.53139 | 0.5422
| 3810000 | 1016 | 2024-11-19 18:28 | 0.62225 | 0 | 0.54005 | 0.0021767 |  | 0.6299 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.53134
| 1000 | 1000 | 2024-11-19 18:28 | 0.6355 | 0 | 0.53348 | 0.0021759 | 1 | #N/A | 0.53133 | 0.53132 | 0.5313 | 0.53128 | 0.54219
| 3840000 | 992 | 2024-11-19 18:29 | 0.62126 | 0 | 0.54003 | 0.0021891 |  | 0.62891 | 0.5422 | 0.54221 | 0.54223 | 0.54224 | 0.53127
| 1000 | 1000 | 2024-11-19 18:30 | 0.63558 | 0 | 0.53346 | 0.0021969 | 1 | #N/A | 0.53127 | 0.53127 | 0.53126 | 0.53126 | 0.54225
| 3840000 | 992 | 2024-11-19 18:31 | 0.62127 | 0 | 0.54007 | 0.002198 |  | 0.62898 | 0.54226 | 0.54227 | 0.54227 | 0.54227 | 0.53128
| 1000 | 1000 | 2024-11-19 18:31 | 0.63557 | 0 | 0.53345 | 0.0021978 | 1 | #N/A | 0.53126 | 0.53126 | 0.53125 | 0.53125 | 0.54224
| 3850000 | 1001 | 2024-11-19 18:32 | 0.62093 | 0 | 0.54006 | 0.0022098 |  | 0.62866 | 0.54226 | 0.54227 | 0.54228 | 0.54229 | 0.53123
| 1000 | 1000 | 2024-11-19 18:33 | 0.63559 | 0 | 0.53342 | 0.0022099 | 1 | #N/A | 0.53122 | 0.53121 | 0.5312 | 0.5312 | 0.54226
| 3850000 | 1001 | 2024-11-19 18:33 | 0.62091 | 0 | 0.54007 | 0.0022123 |  | 0.62867 | 0.54227 | 0.54228 | 0.54228 | 0.54229 | 0.53122
| 1000 | 1000 | 2024-11-19 18:34 | 0.63556 | 0 | 0.5334 | 0.0022088 | 1 | #N/A | 0.53121 | 0.53119 | 0.53118 | 0.53118 | 0.54223
| 3875000 | 992 | 2024-11-19 18:35 | 0.62009 | 0 | 0.54002 | 0.0022359 |  | 0.62791 | 0.54225 | 0.54226 | 0.54226 | 0.54227 | 0.53108
| 1000 | 1000 | 2024-11-19 18:36 | 0.63556 | 0 | 0.5333 | 0.0022349 | 1 | #N/A | 0.53107 | 0.53107 | 0.53106 | 0.53105 | 0.54224
| 3875000 | 992 | 2024-11-19 18:36 | 0.62011 | 0 | 0.54003 | 0.0022309 |  | 0.62792 | 0.54225 | 0.54226 | 0.54226 | 0.54227 | 0.5311
| 1000 | 1000 | 2024-11-19 18:37 | 0.63557 | 0 | 0.53331 | 0.0022327 | 1 | #N/A | 0.53109 | 0.53108 | 0.53107 | 0.53106 | 0.54224
| 3900000 | 1001 | 2024-11-19 18:38 | 0.61931 | 0 | 0.53999 | 0.0022696 |  | 0.62727 | 0.54225 | 0.54226 | 0.54226 | 0.54227 | 0.53091
| 1000 | 1000 | 2024-11-19 18:39 | 0.63555 | 0 | 0.53315 | 0.0022698 | 1 | #N/A | 0.5309 | 0.53088 | 0.53087 | 0.53086 | 0.54223
| 3900000 | 1001 | 2024-11-19 18:39 | 0.61931 | 0 | 0.53998 | 0.0022692 |  | 0.62725 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.5309
| 1000 | 1000 | 2024-11-19 18:40 | 0.63556 | 0 | 0.53315 | 0.0022728 | 1 | #N/A | 0.53089 | 0.53088 | 0.53087 | 0.53086 | 0.54224
| 3900000 | 1001 | 2024-11-19 18:41 | 0.61929 | 0 | 0.53999 | 0.0022718 |  | 0.62724 | 0.54225 | 0.54226 | 0.54227 | 0.54227 | 0.5309
| 1000 | 1000 | 2024-11-19 18:42 | 0.63556 | 0 | 0.53315 | 0.0022697 | 1 | #N/A | 0.53089 | 0.53089 | 0.53087 | 0.53087 | 0.54223
| 3924000 | 981 | 2024-11-19 18:42 | 0.61874 | 0 | 0.53996 | 0.0022886 |  | 0.62673 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.5308
| 1000 | 1000 | 2024-11-19 18:43 | 0.63555 | 0 | 0.53307 | 0.0022881 | 1 | #N/A | 0.5308 | 0.53078 | 0.53078 | 0.53077 | 0.54222
| 3936000 | 984 | 2024-11-19 18:44 | 0.61844 | 0 | 0.53994 | 0.0023045 |  | 0.62649 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53073
| 1000 | 1000 | 2024-11-19 18:45 | 0.63556 | 0 | 0.533 | 0.0023071 | 1 | #N/A | 0.53071 | 0.5307 | 0.53069 | 0.53068 | 0.54223
| 3944000 | 986 | 2024-11-19 18:45 | 0.61826 | 0 | 0.53993 | 0.0023143 |  | 0.62634 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.53068
| 1000 | 1000 | 2024-11-19 18:46 | 0.63556 | 0 | 0.53297 | 0.0023166 | 1 | #N/A | 0.53067 | 0.53065 | 0.53065 | 0.53064 | 0.54224
| 3952000 | 988 | 2024-11-19 18:47 | 0.61816 | 0 | 0.53994 | 0.0023191 |  | 0.62627 | 0.54225 | 0.54226 | 0.54226 | 0.54227 | 0.53066
| 1000 | 1000 | 2024-11-19 18:48 | 0.63554 | 0 | 0.53294 | 0.002318 | 1 | #N/A | 0.53064 | 0.53063 | 0.53062 | 0.53061 | 0.54222
| 3964000 | 991 | 2024-11-19 18:48 | 0.61792 | 0 | 0.53991 | 0.0023303 |  | 0.62604 | 0.54223 | 0.54223 | 0.54224 | 0.54225 | 0.53059
| 1000 | 1000 | 2024-11-19 18:49 | 0.63555 | 0 | 0.5329 | 0.0023323 | 1 | #N/A | 0.53058 | 0.53057 | 0.53056 | 0.53055 | 0.54223
| 3972000 | 993 | 2024-11-19 18:50 | 0.61772 | 0 | 0.53989 | 0.0023499 |  | 0.62593 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.53049
| 1000 | 1000 | 2024-11-19 18:51 | 0.63555 | 0 | 0.53281 | 0.0023523 | 1 | #N/A | 0.53048 | 0.53047 | 0.53046 | 0.53045 | 0.54222
| 3980000 | 995 | 2024-11-19 18:51 | 0.61764 | 0 | 0.53989 | 0.0023557 |  | 0.62587 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.53047
| 1000 | 1000 | 2024-11-19 18:52 | 0.63554 | 0 | 0.53279 | 0.0023558 | 1 | #N/A | 0.53045 | 0.53044 | 0.53043 | 0.53042 | 0.54222
| 3992000 | 998 | 2024-11-19 18:53 | 0.61751 | 0 | 0.53988 | 0.0023633 |  | 0.62575 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.53042
| 1000 | 1000 | 2024-11-19 18:53 | 0.63554 | 0 | 0.53276 | 0.0023636 | 1 | #N/A | 0.53041 | 0.5304 | 0.53039 | 0.53039 | 0.54222
| 4000000 | 1000 | 2024-11-19 18:54 | 0.61741 | 0 | 0.53987 | 0.0023742 |  | 0.6257 | 0.54223 | 0.54223 | 0.54224 | 0.54225 | 0.53037
| 1000 | 1000 | 2024-11-19 18:55 | 0.63553 | 0 | 0.53271 | 0.0023746 | 1 | #N/A | 0.53035 | 0.53034 | 0.53033 | 0.53032 | 0.54221
| 4008000 | 1002 | 2024-11-19 18:56 | 0.61735 | 0 | 0.53986 | 0.0023807 |  | 0.62565 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53033
| 1000 | 1000 | 2024-11-19 18:56 | 0.63554 | 0 | 0.53269 | 0.002382 | 1 | #N/A | 0.53032 | 0.53031 | 0.5303 | 0.5303 | 0.54222
| 4020000 | 1005 | 2024-11-19 18:57 | 0.6173 | 0 | 0.53985 | 0.0023907 |  | 0.62563 | 0.54223 | 0.54223 | 0.54225 | 0.54226 | 0.53029
| 1000 | 1000 | 2024-11-19 18:58 | 0.63556 | 0 | 0.53266 | 0.0023926 | 1 | #N/A | 0.53028 | 0.53027 | 0.53027 | 0.53026 | 0.54223
| 4028000 | 1007 | 2024-11-19 18:59 | 0.61723 | 0 | 0.53985 | 0.0024016 |  | 0.62562 | 0.54224 | 0.54224 | 0.54226 | 0.54227 | 0.53024
| 1000 | 1000 | 2024-11-19 18:59 | 0.63552 | 0 | 0.53261 | 0.0023982 | 1 | #N/A | 0.53023 | 0.53022 | 0.53021 | 0.5302 | 0.5422
| 4036000 | 1009 | 2024-11-19 19:00 | 0.61718 | 0 | 0.53982 | 0.002414 |  | 0.6256 | 0.54222 | 0.54223 | 0.54223 | 0.54225 | 0.53016
| 1000 | 1000 | 2024-11-19 19:01 | 0.63553 | 0 | 0.53255 | 0.0024145 | 1 | #N/A | 0.53015 | 0.53014 | 0.53013 | 0.53012 | 0.54221
| 4050000 | 999 | 2024-11-19 19:02 | 0.61712 | 0 | 0.5398 | 0.0024287 |  | 0.62559 | 0.54222 | 0.54222 | 0.54224 | 0.54224 | 0.53009
| 1000 | 1000 | 2024-11-19 19:02 | 0.63553 | 0 | 0.53249 | 0.0024285 | 1 | #N/A | 0.53008 | 0.53007 | 0.53006 | 0.53005 | 0.54221
| 4050000 | 999 | 2024-11-19 19:03 | 0.61715 | 0 | 0.5398 | 0.0024283 |  | 0.62562 | 0.54221 | 0.54222 | 0.54224 | 0.54225 | 0.53009
| 1000 | 1000 | 2024-11-19 19:04 | 0.63554 | 0 | 0.53249 | 0.0024312 | 1 | #N/A | 0.53007 | 0.53007 | 0.53006 | 0.53005 | 0.54222
| 4064000 | 1016 | 2024-11-19 19:04 | 0.61718 | 0 | 0.5398 | 0.0024395 |  | 0.6257 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.53004
| 1000 | 1000 | 2024-11-19 19:05 | 0.63554 | 0 | 0.53246 | 0.00244 | 1 | #N/A | 0.53003 | 0.53002 | 0.53001 | 0.53 | 0.54222
| 4080000 | 986 | 2024-11-19 19:06 | 0.61721 | 0 | 0.53978 | 0.0024572 |  | 0.62578 | 0.54223 | 0.54223 | 0.54224 | 0.54226 | 0.52995
| 1000 | 1000 | 2024-11-19 19:07 | 0.63555 | 0 | 0.53239 | 0.0024584 | 1 | #N/A | 0.52994 | 0.52993 | 0.52992 | 0.52992 | 0.54222
| 4080000 | 986 | 2024-11-19 19:07 | 0.61721 | 0 | 0.53978 | 0.002459 |  | 0.62579 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.52994
| 1000 | 1000 | 2024-11-19 19:08 | 0.63555 | 0 | 0.53238 | 0.0024603 | 1 | #N/A | 0.52993 | 0.52993 | 0.52992 | 0.52992 | 0.54222
| 4100000 | 984 | 2024-11-19 19:09 | 0.61734 | 0 | 0.53977 | 0.0024852 |  | 0.62603 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.52983
| 1000 | 1000 | 2024-11-19 19:10 | 0.63555 | 0 | 0.53228 | 0.0024838 | 1 | #N/A | 0.52981 | 0.5298 | 0.52979 | 0.52978 | 0.54222
| 4125000 | 1001 | 2024-11-19 19:10 | 0.61763 | 0 | 0.53973 | 0.0025082 |  | 0.62639 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.5297
| 1000 | 1000 | 2024-11-19 19:11 | 0.63555 | 0 | 0.53218 | 0.0025104 | 1 | #N/A | 0.52969 | 0.52968 | 0.52967 | 0.52966 | 0.54222
| 4125000 | 1001 | 2024-11-19 19:12 | 0.6176 | 0 | 0.53974 | 0.0025111 |  | 0.62637 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.5297
| 1000 | 1000 | 2024-11-19 19:13 | 0.63555 | 0 | 0.53219 | 0.002509 | 1 | #N/A | 0.52969 | 0.52968 | 0.52967 | 0.52967 | 0.54222
| 4125000 | 1001 | 2024-11-19 19:13 | 0.61762 | 0 | 0.53974 | 0.0025098 |  | 0.62639 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.5297
| 1000 | 1000 | 2024-11-19 19:14 | 0.63556 | 0 | 0.53218 | 0.0025112 | 1 | #N/A | 0.52969 | 0.52968 | 0.52966 | 0.52965 | 0.54223
| 4150000 | 996 | 2024-11-19 19:15 | 0.61797 | 0 | 0.5397 | 0.0025393 |  | 0.62685 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.52955
| 1000 | 1000 | 2024-11-19 19:16 | 0.63554 | 0 | 0.53206 | 0.002539 | 1 | #N/A | 0.52954 | 0.52953 | 0.52952 | 0.52951 | 0.54222
| 4150000 | 996 | 2024-11-19 19:16 | 0.61799 | 0 | 0.5397 | 0.0025377 |  | 0.62686 | 0.54223 | 0.54223 | 0.54225 | 0.54225 | 0.52955
| 1000 | 1000 | 2024-11-19 19:17 | 0.63555 | 0 | 0.53206 | 0.00254 | 1 | #N/A | 0.52953 | 0.52952 | 0.52952 | 0.52952 | 0.54222
| 4150000 | 996 | 2024-11-19 19:18 | 0.61797 | 0 | 0.5397 | 0.0025396 |  | 0.62684 | 0.54223 | 0.54224 | 0.54224 | 0.54226 | 0.52954
| 1000 | 1000 | 2024-11-19 19:19 | 0.63555 | 0 | 0.53206 | 0.0025405 | 1 | #N/A | 0.52953 | 0.52952 | 0.52951 | 0.5295 | 0.54222
| 4175000 | 1002 | 2024-11-19 19:19 | 0.61844 | 0 | 0.53967 | 0.0025684 |  | 0.62743 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.5294
| 1000 | 1000 | 2024-11-19 19:20 | 0.63554 | 0 | 0.53194 | 0.0025673 | 1 | #N/A | 0.52939 | 0.52938 | 0.52937 | 0.52935 | 0.54221
| 4175000 | 1002 | 2024-11-19 19:21 | 0.6184 | 0 | 0.53966 | 0.0025713 |  | 0.62741 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.52938
| 1000 | 1000 | 2024-11-19 19:22 | 0.63553 | 0 | 0.53192 | 0.0025718 | 1 | #N/A | 0.52936 | 0.52935 | 0.52934 | 0.52933 | 0.54221
| 4200000 | 1001 | 2024-11-19 19:22 | 0.61906 | 0 | 0.53965 | 0.0025863 |  | 0.62812 | 0.54222 | 0.54223 | 0.54224 | 0.54226 | 0.52931
| 1000 | 1000 | 2024-11-19 19:23 | 0.63556 | 0 | 0.53187 | 0.0025904 | 1 | #N/A | 0.52929 | 0.52928 | 0.52927 | 0.52926 | 0.54223
| 4200000 | 1001 | 2024-11-19 19:24 | 0.61905 | 0 | 0.53967 | 0.0025898 |  | 0.62813 | 0.54224 | 0.54225 | 0.54227 | 0.54228 | 0.52931
| 1000 | 1000 | 2024-11-19 19:24 | 0.63556 | 0 | 0.53187 | 0.0025883 | 1 | #N/A | 0.5293 | 0.52929 | 0.52928 | 0.52927 | 0.54223
| 4200000 | 1001 | 2024-11-19 19:25 | 0.61904 | 0 | 0.53966 | 0.0025887 |  | 0.6281 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.52931
| 1000 | 1000 | 2024-11-19 19:26 | 0.63558 | 0 | 0.53188 | 0.0025909 | 1 | #N/A | 0.5293 | 0.52929 | 0.52928 | 0.52928 | 0.54224
| 4225000 | 1014 | 2024-11-19 19:27 | 0.61963 | 0 | 0.53963 | 0.0026274 |  | 0.62885 | 0.54225 | 0.54226 | 0.54227 | 0.54227 | 0.52912
| 1000 | 1000 | 2024-11-19 19:27 | 0.63555 | 0 | 0.53172 | 0.0026247 | 1 | #N/A | 0.52911 | 0.5291 | 0.52909 | 0.52908 | 0.54222
| 4230000 | 987 | 2024-11-19 19:28 | 0.6198 | 0 | 0.53962 | 0.0026238 |  | 0.62901 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.52912
| 1000 | 1000 | 2024-11-19 19:29 | 0.63555 | 0 | 0.53172 | 0.0026258 | 1 | #N/A | 0.52911 | 0.52909 | 0.52908 | 0.52907 | 0.54222
| 4250000 | 1003 | 2024-11-19 19:30 | 0.62036 | 0 | 0.5396 | 0.0026502 |  | 0.62965 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.529
| 1000 | 1000 | 2024-11-19 19:30 | 0.63555 | 0 | 0.53163 | 0.0026499 | 1 | #N/A | 0.52899 | 0.52898 | 0.52897 | 0.52897 | 0.54223
| 4250000 | 1003 | 2024-11-19 19:31 | 0.62037 | 0 | 0.5396 | 0.0026491 |  | 0.62967 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.529
| 1000 | 1000 | 2024-11-19 19:32 | 0.63557 | 0 | 0.53163 | 0.0026523 | 1 | #N/A | 0.52899 | 0.52898 | 0.52897 | 0.52896 | 0.54224
| 4260000 | 994 | 2024-11-19 19:33 | 0.62066 | 0 | 0.53959 | 0.0026634 |  | 0.63002 | 0.54225 | 0.54226 | 0.54226 | 0.54227 | 0.52894
| 1000 | 1000 | 2024-11-19 19:33 | 0.63556 | 0 | 0.53158 | 0.0026634 | 1 | #N/A | 0.52893 | 0.52892 | 0.52891 | 0.5289 | 0.54223
| 4260000 | 994 | 2024-11-19 19:34 | 0.62066 | 0 | 0.53959 | 0.0026643 |  | 0.63002 | 0.54224 | 0.54225 | 0.54227 | 0.54227 | 0.52893
| 1000 | 1000 | 2024-11-19 19:35 | 0.63557 | 0 | 0.53158 | 0.0026648 | 1 | #N/A | 0.52892 | 0.52892 | 0.52891 | 0.5289 | 0.54224
| 4290000 | 1001 | 2024-11-19 19:36 | 0.62161 | 0 | 0.53957 | 0.0026975 |  | 0.63112 | 0.54225 | 0.54226 | 0.54227 | 0.54228 | 0.52878
| 1000 | 1000 | 2024-11-19 19:36 | 0.63553 | 0 | 0.53144 | 0.0026935 | 1 | #N/A | 0.52876 | 0.52875 | 0.52874 | 0.52873 | 0.54221
| 4290000 | 1001 | 2024-11-19 19:37 | 0.62161 | 0 | 0.53954 | 0.0026939 |  | 0.63108 | 0.54222 | 0.54222 | 0.54223 | 0.54225 | 0.52876
| 1000 | 1000 | 2024-11-19 19:38 | 0.63554 | 0 | 0.53144 | 0.002695 | 1 | #N/A | 0.52875 | 0.52875 | 0.52874 | 0.52872 | 0.54222
| 4290000 | 1001 | 2024-11-19 19:38 | 0.62162 | 0 | 0.53955 | 0.0026953 |  | 0.63111 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.52877
| 1000 | 1000 | 2024-11-19 19:39 | 0.63555 | 0 | 0.53144 | 0.0026963 | 1 | #N/A | 0.52876 | 0.52875 | 0.52873 | 0.52873 | 0.54222
| 4320000 | 1008 | 2024-11-19 19:40 | 0.62267 | 0 | 0.53952 | 0.0027227 |  | 0.63227 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.52863
| 1000 | 1000 | 2024-11-19 19:41 | 0.63556 | 0 | 0.53133 | 0.0027253 | 1 | #N/A | 0.52862 | 0.52861 | 0.5286 | 0.52859 | 0.54224
| 4320000 | 1008 | 2024-11-19 19:41 | 0.62266 | 0 | 0.53953 | 0.0027245 |  | 0.63226 | 0.54225 | 0.54225 | 0.54226 | 0.54227 | 0.52863
| 1000 | 1000 | 2024-11-19 19:42 | 0.63557 | 0 | 0.53134 | 0.002726 | 1 | #N/A | 0.52862 | 0.52862 | 0.52861 | 0.52861 | 0.54225
| 4320000 | 1008 | 2024-11-19 19:43 | 0.62266 | 0 | 0.53954 | 0.0027256 |  | 0.63226 | 0.54225 | 0.54226 | 0.54227 | 0.54227 | 0.52863
| 1000 | 1000 | 2024-11-19 19:44 | 0.63556 | 0 | 0.53134 | 0.0027241 | 1 | #N/A | 0.52862 | 0.52862 | 0.52861 | 0.52861 | 0.54223
| 4350000 | 986 | 2024-11-19 19:44 | 0.62365 | 0 | 0.53949 | 0.0027649 |  | 0.63342 | 0.54225 | 0.54225 | 0.54226 | 0.54226 | 0.52843
| 1000 | 1000 | 2024-11-19 19:45 | 0.63556 | 0 | 0.53117 | 0.0027645 | 1 | #N/A | 0.52842 | 0.52841 | 0.5284 | 0.5284 | 0.54223
| 4350000 | 986 | 2024-11-19 19:46 | 0.62366 | 0 | 0.53948 | 0.0027615 |  | 0.63342 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.52844
| 1000 | 1000 | 2024-11-19 19:47 | 0.63555 | 0 | 0.53117 | 0.0027634 | 1 | #N/A | 0.52842 | 0.52841 | 0.5284 | 0.52839 | 0.54222
| 4375000 | 1001 | 2024-11-19 19:47 | 0.62449 | 0 | 0.53945 | 0.0027965 |  | 0.63438 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.52826
| 1000 | 1000 | 2024-11-19 19:48 | 0.63555 | 0 | 0.53104 | 0.0027966 | 1 | #N/A | 0.52826 | 0.52824 | 0.52824 | 0.52823 | 0.54222
| 4375000 | 1001 | 2024-11-19 19:49 | 0.6245 | 0 | 0.53946 | 0.0027931 |  | 0.63439 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.52828
| 1000 | 1000 | 2024-11-19 19:50 | 0.63555 | 0 | 0.53105 | 0.0027913 | 1 | #N/A | 0.52827 | 0.52826 | 0.52826 | 0.52825 | 0.54222
| 4375000 | 1001 | 2024-11-19 19:50 | 0.62449 | 0 | 0.53945 | 0.0027926 |  | 0.63437 | 0.54222 | 0.54224 | 0.54225 | 0.54225 | 0.52828
| 1000 | 1000 | 2024-11-19 19:51 | 0.63554 | 0 | 0.53104 | 0.0027944 | 1 | #N/A | 0.52826 | 0.52825 | 0.52824 | 0.52823 | 0.54222
| 4400000 | 990 | 2024-11-19 19:52 | 0.62543 | 0 | 0.53943 | 0.0028066 |  | 0.63535 | 0.54223 | 0.54223 | 0.54225 | 0.54226 | 0.52821
| 1000 | 1000 | 2024-11-19 19:53 | 0.6356 | 0 | 0.53101 | 0.0028118 | 1 | #N/A | 0.5282 | 0.52821 | 0.5282 | 0.5282 | 0.54226
| 4400000 | 990 | 2024-11-19 19:53 | 0.62547 | 0 | 0.53947 | 0.0028088 |  | 0.63542 | 0.54227 | 0.54227 | 0.54227 | 0.54228 | 0.52823
| 1000 | 1000 | 2024-11-19 19:54 | 0.6356 | 0 | 0.53102 | 0.0028112 | 1 | #N/A | 0.52822 | 0.52821 | 0.5282 | 0.5282 | 0.54226
| 4400000 | 990 | 2024-11-19 19:55 | 0.62545 | 0 | 0.53947 | 0.0028105 |  | 0.63539 | 0.54227 | 0.54228 | 0.54228 | 0.54228 | 0.52822
| 1000 | 1000 | 2024-11-19 19:56 | 0.63561 | 0 | 0.53102 | 0.0028119 | 1 | #N/A | 0.52822 | 0.52822 | 0.52821 | 0.5282 | 0.54227
| 4425000 | 1003 | 2024-11-19 19:56 | 0.62617 | 0 | 0.53942 | 0.0028599 |  | 0.63633 | 0.54228 | 0.54228 | 0.54228 | 0.54229 | 0.52798
| 1000 | 1000 | 2024-11-19 19:57 | 0.63557 | 0 | 0.53081 | 0.0028574 | 1 | #N/A | 0.52797 | 0.52795 | 0.52794 | 0.52793 | 0.54224
| 4425000 | 1003 | 2024-11-19 19:58 | 0.62615 | 0 | 0.5394 | 0.0028582 |  | 0.63631 | 0.54225 | 0.54226 | 0.54226 | 0.54227 | 0.52797
| 1000 | 1000 | 2024-11-19 19:58 | 0.63553 | 0 | 0.53079 | 0.0028548 | 1 | #N/A | 0.52795 | 0.52794 | 0.52792 | 0.52791 | 0.54221
| 4440000 | 999 | 2024-11-19 19:59 | 0.62663 | 0 | 0.53936 | 0.0028657 |  | 0.63679 | 0.54222 | 0.54222 | 0.54223 | 0.54224 | 0.5279
| 1000 | 1000 | 2024-11-19 20:00 | 0.63557 | 0 | 0.53076 | 0.00287 | 1 | #N/A | 0.52789 | 0.52789 | 0.52789 | 0.52788 | 0.54224
| 4440000 | 999 | 2024-11-19 20:01 | 0.62666 | 0 | 0.53939 | 0.0028674 |  | 0.63685 | 0.54225 | 0.54225 | 0.54226 | 0.54227 | 0.52792
| 1000 | 1000 | 2024-11-19 20:01 | 0.63554 | 0 | 0.53076 | 0.0028637 | 1 | #N/A | 0.52791 | 0.5279 | 0.52789 | 0.52788 | 0.54221
| 4440000 | 999 | 2024-11-19 20:02 | 0.62666 | 0 | 0.53938 | 0.0028643 |  | 0.63683 | 0.54222 | 0.54224 | 0.54224 | 0.54226 | 0.52792
| 1000 | 1000 | 2024-11-19 20:03 | 0.63556 | 0 | 0.53076 | 0.0028675 | 1 | #N/A | 0.52791 | 0.5279 | 0.52789 | 0.52788 | 0.54223
| 4480000 | 1008 | 2024-11-19 20:04 | 0.62784 | 0 | 0.53934 | 0.0029104 |  | 0.63822 | 0.54224 | 0.54225 | 0.54225 | 0.54227 | 0.5277
| 1000 | 1000 | 2024-11-19 20:04 | 0.63554 | 0 | 0.53058 | 0.0029092 | 1 | #N/A | 0.52769 | 0.52767 | 0.52766 | 0.52765 | 0.54221
| 4480000 | 1008 | 2024-11-19 20:05 | 0.62783 | 0 | 0.53933 | 0.002911 |  | 0.63819 | 0.54222 | 0.54224 | 0.54225 | 0.54227 | 0.52769
| 1000 | 1000 | 2024-11-19 20:06 | 0.63556 | 0 | 0.53058 | 0.0029142 | 1 | #N/A | 0.52768 | 0.52767 | 0.52766 | 0.52765 | 0.54224
| 4500000 | 999 | 2024-11-19 20:07 | 0.62832 | 0 | 0.53932 | 0.0029448 |  | 0.63882 | 0.54225 | 0.54226 | 0.54227 | 0.54228 | 0.52754
| 1000 | 1000 | 2024-11-19 20:07 | 0.63556 | 0 | 0.53046 | 0.0029449 | 1 | #N/A | 0.52753 | 0.52752 | 0.52751 | 0.5275 | 0.54224
| 4500000 | 999 | 2024-11-19 20:08 | 0.62832 | 0 | 0.53931 | 0.0029416 |  | 0.6388 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.52755
| 1000 | 1000 | 2024-11-19 20:09 | 0.63556 | 0 | 0.53046 | 0.0029409 | 1 | #N/A | 0.52754 | 0.52752 | 0.52752 | 0.52752 | 0.54223
| 4500000 | 999 | 2024-11-19 20:09 | 0.62832 | 0 | 0.5393 | 0.0029411 |  | 0.63881 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.52754
| 1000 | 1000 | 2024-11-19 20:10 | 0.63556 | 0 | 0.53045 | 0.0029431 | 1 | #N/A | 0.52753 | 0.52752 | 0.52751 | 0.52749 | 0.54223
| 4520000 | 1017 | 2024-11-19 20:11 | 0.62877 | 0 | 0.53929 | 0.0029663 |  | 0.63936 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.52742
| 1000 | 1000 | 2024-11-19 20:12 | 0.63555 | 0 | 0.53036 | 0.0029639 | 1 | #N/A | 0.52741 | 0.5274 | 0.52739 | 0.52738 | 0.54222
| 4520000 | 1017 | 2024-11-19 20:12 | 0.62879 | 0 | 0.53928 | 0.0029656 |  | 0.63937 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.52742
| 1000 | 1000 | 2024-11-19 20:13 | 0.63554 | 0 | 0.53036 | 0.0029641 | 1 | #N/A | 0.52741 | 0.5274 | 0.52739 | 0.52738 | 0.54222
| 4550000 | 1001 | 2024-11-19 20:14 | 0.62938 | 0 | 0.53925 | 0.003 |  | 0.6401 | 0.54223 | 0.54224 | 0.54225 | 0.54227 | 0.52725
| 1000 | 1000 | 2024-11-19 20:15 | 0.63554 | 0 | 0.53022 | 0.002999 | 1 | #N/A | 0.52723 | 0.52722 | 0.52721 | 0.5272 | 0.54221
| 4550000 | 1001 | 2024-11-19 20:15 | 0.62936 | 0 | 0.53923 | 0.0030008 |  | 0.64008 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.52723
| 1000 | 1000 | 2024-11-19 20:16 | 0.63553 | 0 | 0.5302 | 0.0030012 | 1 | #N/A | 0.52722 | 0.52721 | 0.5272 | 0.52719 | 0.54221
| 4550000 | 1001 | 2024-11-19 20:17 | 0.62939 | 0 | 0.53923 | 0.0029989 |  | 0.6401 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.52723
| 1000 | 1000 | 2024-11-19 20:18 | 0.63553 | 0 | 0.53021 | 0.003 | 1 | #N/A | 0.52722 | 0.52721 | 0.5272 | 0.52719 | 0.54221
| 4560000 | 988 | 2024-11-19 20:18 | 0.62955 | 0 | 0.53922 | 0.0030094 |  | 0.64031 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.52718
| 1000 | 1000 | 2024-11-19 20:19 | 0.63553 | 0 | 0.53016 | 0.0030107 | 1 | #N/A | 0.52717 | 0.52716 | 0.52714 | 0.52714 | 0.54221
| 4600000 | 989 | 2024-11-19 20:20 | 0.63003 | 0 | 0.53918 | 0.0030573 |  | 0.64095 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.52695
| 1000 | 1000 | 2024-11-19 20:21 | 0.63554 | 0 | 0.52998 | 0.0030578 | 1 | #N/A | 0.52694 | 0.52693 | 0.52692 | 0.52691 | 0.54222
| 4600000 | 989 | 2024-11-19 20:21 | 0.63005 | 0 | 0.53919 | 0.0030541 |  | 0.64098 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.52697
| 1000 | 1000 | 2024-11-19 20:22 | 0.63555 | 0 | 0.53 | 0.0030572 | 1 | #N/A | 0.52696 | 0.52694 | 0.52693 | 0.52692 | 0.54222
| 4600000 | 989 | 2024-11-19 20:23 | 0.63005 | 0 | 0.53919 | 0.0030577 |  | 0.64098 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.52696
| 1000 | 1000 | 2024-11-19 20:24 | 0.63555 | 0 | 0.52999 | 0.0030576 | 1 | #N/A | 0.52695 | 0.52694 | 0.52693 | 0.52693 | 0.54222
| 4620000 | 1001 | 2024-11-19 20:24 | 0.63016 | 0 | 0.53916 | 0.0030842 |  | 0.64119 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.52682
| 1000 | 1000 | 2024-11-19 20:25 | 0.63554 | 0 | 0.52989 | 0.0030827 | 1 | #N/A | 0.52682 | 0.52681 | 0.5268 | 0.52679 | 0.54222
| 4620000 | 1001 | 2024-11-19 20:26 | 0.63015 | 0 | 0.53915 | 0.0030851 |  | 0.64121 | 0.54222 | 0.54224 | 0.54224 | 0.54225 | 0.52681
| 1000 | 1000 | 2024-11-19 20:27 | 0.63553 | 0 | 0.52986 | 0.003086 | 1 | #N/A | 0.52679 | 0.52678 | 0.52677 | 0.52676 | 0.5422
| 4625000 | 999 | 2024-11-19 20:27 | 0.63016 | 0 | 0.53914 | 0.0030952 |  | 0.64124 | 0.54221 | 0.54223 | 0.54224 | 0.54225 | 0.52675
| 1000 | 1000 | 2024-11-19 20:28 | 0.63553 | 0 | 0.52982 | 0.003095 | 1 | #N/A | 0.52674 | 0.52673 | 0.52672 | 0.52671 | 0.5422
| 4650000 | 992 | 2024-11-19 20:29 | 0.6302 | 0 | 0.5391 | 0.003122 |  | 0.64138 | 0.54221 | 0.54222 | 0.54223 | 0.54223 | 0.52661
| 1000 | 1000 | 2024-11-19 20:29 | 0.63552 | 0 | 0.5297 | 0.0031236 | 1 | #N/A | 0.5266 | 0.52658 | 0.52657 | 0.52656 | 0.5422
| 4650000 | 992 | 2024-11-19 20:30 | 0.6302 | 0 | 0.53911 | 0.0031211 |  | 0.64137 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.52662
| 1000 | 1000 | 2024-11-19 20:31 | 0.63554 | 0 | 0.52972 | 0.0031242 | 1 | #N/A | 0.52661 | 0.5266 | 0.52659 | 0.52658 | 0.54221
| 4650000 | 992 | 2024-11-19 20:32 | 0.63019 | 0 | 0.53911 | 0.0031244 |  | 0.64137 | 0.54222 | 0.54224 | 0.54224 | 0.54225 | 0.52661
| 1000 | 1000 | 2024-11-19 20:32 | 0.63554 | 0 | 0.52971 | 0.0031261 | 1 | #N/A | 0.5266 | 0.52659 | 0.52658 | 0.52658 | 0.54222
| 4680000 | 1014 | 2024-11-19 20:33 | 0.63011 | 0 | 0.53909 | 0.0031567 |  | 0.64142 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.52646
| 1000 | 1000 | 2024-11-19 20:34 | 0.63553 | 0 | 0.52958 | 0.003155 | 1 | #N/A | 0.52645 | 0.52643 | 0.52642 | 0.52641 | 0.5422
| 4680000 | 1014 | 2024-11-19 20:35 | 0.63008 | 0 | 0.53908 | 0.0031563 |  | 0.64138 | 0.54221 | 0.54223 | 0.54224 | 0.54225 | 0.52645
| 1000 | 1000 | 2024-11-19 20:35 | 0.63553 | 0 | 0.52958 | 0.0031557 | 1 | #N/A | 0.52644 | 0.52643 | 0.52642 | 0.52641 | 0.54221
| 4700000 | 987 | 2024-11-19 20:36 | 0.62984 | 0 | 0.53905 | 0.0031867 |  | 0.64125 | 0.54222 | 0.54224 | 0.54224 | 0.54225 | 0.5263
| 1000 | 1000 | 2024-11-19 20:37 | 0.63552 | 0 | 0.52946 | 0.0031852 | 1 | #N/A | 0.52629 | 0.52628 | 0.52627 | 0.52626 | 0.5422
| 4700000 | 987 | 2024-11-19 20:38 | 0.62983 | 0 | 0.53905 | 0.0031865 |  | 0.64124 | 0.54222 | 0.54224 | 0.54225 | 0.54225 | 0.52631
| 1000 | 1000 | 2024-11-19 20:38 | 0.63553 | 0 | 0.52946 | 0.0031852 | 1 | #N/A | 0.5263 | 0.52628 | 0.52627 | 0.52626 | 0.5422
| 4700000 | 987 | 2024-11-19 20:39 | 0.62986 | 0 | 0.53904 | 0.0031853 |  | 0.64126 | 0.54221 | 0.54223 | 0.54223 | 0.54224 | 0.5263
| 1000 | 1000 | 2024-11-19 20:40 | 0.63553 | 0 | 0.52947 | 0.003185 | 1 | #N/A | 0.52629 | 0.52629 | 0.52628 | 0.52627 | 0.54221
| 4725000 | 1008 | 2024-11-19 20:41 | 0.62956 | 0 | 0.53902 | 0.0032095 |  | 0.64104 | 0.54222 | 0.54222 | 0.54223 | 0.54224 | 0.52618
| 1000 | 1000 | 2024-11-19 20:41 | 0.63553 | 0 | 0.52936 | 0.0032101 | 1 | #N/A | 0.52617 | 0.52616 | 0.52615 | 0.52614 | 0.54221
| 4725000 | 1008 | 2024-11-19 20:42 | 0.62957 | 0 | 0.53902 | 0.0032108 |  | 0.64105 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.52618
| 1000 | 1000 | 2024-11-19 20:43 | 0.63554 | 0 | 0.52937 | 0.0032102 | 1 | #N/A | 0.52617 | 0.52617 | 0.52616 | 0.52615 | 0.54221
| 4750000 | 1007 | 2024-11-19 20:43 | 0.62911 | 0 | 0.539 | 0.0032384 |  | 0.64067 | 0.54222 | 0.54224 | 0.54224 | 0.54226 | 0.52605
| 1000 | 1000 | 2024-11-19 20:44 | 0.63556 | 0 | 0.52927 | 0.00324 | 1 | #N/A | 0.52604 | 0.52604 | 0.52603 | 0.52602 | 0.54223
| 4750000 | 1007 | 2024-11-19 20:45 | 0.62913 | 0 | 0.53903 | 0.0032398 |  | 0.64071 | 0.54225 | 0.54226 | 0.54227 | 0.54229 | 0.52607
| 1000 | 1000 | 2024-11-19 20:46 | 0.63555 | 0 | 0.52928 | 0.0032361 | 1 | #N/A | 0.52606 | 0.52605 | 0.52605 | 0.52604 | 0.54223
| 4750000 | 1007 | 2024-11-19 20:46 | 0.62916 | 0 | 0.53902 | 0.0032356 |  | 0.64074 | 0.54224 | 0.54224 | 0.54226 | 0.54227 | 0.52608
| 1000 | 1000 | 2024-11-19 20:47 | 0.63555 | 0 | 0.52928 | 0.0032357 | 1 | #N/A | 0.52606 | 0.52605 | 0.52604 | 0.52603 | 0.54222
| 4776000 | 995 | 2024-11-19 20:48 | 0.62854 | 0 | 0.53898 | 0.0032671 |  | 0.64019 | 0.54223 | 0.54224 | 0.54226 | 0.54227 | 0.52591
| 1000 | 1000 | 2024-11-19 20:49 | 0.63557 | 0 | 0.52917 | 0.0032667 | 1 | #N/A | 0.52591 | 0.5259 | 0.5259 | 0.52589 | 0.54223
| 4776000 | 995 | 2024-11-19 20:49 | 0.62853 | 0 | 0.53899 | 0.0032674 |  | 0.6402 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.52592
| 1000 | 1000 | 2024-11-19 20:50 | 0.63558 | 0 | 0.52916 | 0.0032712 | 1 | #N/A | 0.52591 | 0.5259 | 0.52588 | 0.52588 | 0.54225
| 4800000 | 1000 | 2024-11-19 20:51 | 0.62784 | 0 | 0.53897 | 0.0033037 |  | 0.63963 | 0.54226 | 0.54227 | 0.54227 | 0.54228 | 0.52575
| 1000 | 1000 | 2024-11-19 20:52 | 0.63556 | 0 | 0.52903 | 0.0032988 | 1 | #N/A | 0.52574 | 0.52573 | 0.52573 | 0.52573 | 0.54223
| 4800000 | 1000 | 2024-11-19 20:52 | 0.62779 | 0 | 0.53894 | 0.0033092 |  | 0.63961 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.52571
| 1000 | 1000 | 2024-11-19 20:53 | 0.63554 | 0 | 0.52899 | 0.0033054 | 1 | #N/A | 0.52569 | 0.52568 | 0.52568 | 0.52566 | 0.54221
| 4800000 | 1000 | 2024-11-19 20:54 | 0.62784 | 0 | 0.53893 | 0.0032948 |  | 0.6396 | 0.54222 | 0.54222 | 0.54224 | 0.54225 | 0.52576
| 1000 | 1000 | 2024-11-19 20:55 | 0.63555 | 0 | 0.52903 | 0.0032982 | 1 | #N/A | 0.52574 | 0.52573 | 0.52572 | 0.52572 | 0.54222
| 4824000 | 1005 | 2024-11-19 20:55 | 0.62703 | 0 | 0.53891 | 0.0033307 |  | 0.6389 | 0.54223 | 0.54223 | 0.54224 | 0.54225 | 0.52558
| 1000 | 1000 | 2024-11-19 20:56 | 0.63555 | 0 | 0.52889 | 0.0033323 | 1 | #N/A | 0.52557 | 0.52556 | 0.52556 | 0.52554 | 0.54222
| 4824000 | 1005 | 2024-11-19 20:57 | 0.627 | 0 | 0.53891 | 0.0033342 |  | 0.63889 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.52557
| 1000 | 1000 | 2024-11-19 20:58 | 0.63554 | 0 | 0.52888 | 0.0033312 | 1 | #N/A | 0.52556 | 0.52556 | 0.52555 | 0.52553 | 0.54221
| 4848000 | 1010 | 2024-11-19 20:58 | 0.62614 | 0 | 0.53887 | 0.0033606 |  | 0.6381 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.52543
| 1000 | 1000 | 2024-11-19 20:59 | 0.63554 | 0 | 0.52877 | 0.0033618 | 1 | #N/A | 0.52542 | 0.52541 | 0.5254 | 0.5254 | 0.54221
| 4848000 | 1010 | 2024-11-19 21:00 | 0.62615 | 0 | 0.53888 | 0.0033609 |  | 0.63812 | 0.54222 | 0.54223 | 0.54225 | 0.54226 | 0.52544
| 1000 | 1000 | 2024-11-19 21:00 | 0.63555 | 0 | 0.52877 | 0.0033622 | 1 | #N/A | 0.52543 | 0.52542 | 0.5254 | 0.52539 | 0.54222
| 4875000 | 1001 | 2024-11-19 21:01 | 0.62506 | 0 | 0.53883 | 0.0034003 |  | 0.63716 | 0.54223 | 0.54223 | 0.54224 | 0.54224 | 0.52523
| 1000 | 1000 | 2024-11-19 21:02 | 0.63554 | 0 | 0.52861 | 0.0034025 | 1 | #N/A | 0.52522 | 0.52521 | 0.5252 | 0.52519 | 0.54222
| 4875000 | 1001 | 2024-11-19 21:03 | 0.62506 | 0 | 0.53883 | 0.0034007 |  | 0.63716 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.52523
| 1000 | 1000 | 2024-11-19 21:03 | 0.63552 | 0 | 0.5286 | 0.0033996 | 1 | #N/A | 0.52522 | 0.52521 | 0.5252 | 0.52519 | 0.5422
| 4875000 | 1001 | 2024-11-19 21:04 | 0.62503 | 0 | 0.53883 | 0.0033992 |  | 0.63713 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.52523
| 1000 | 1000 | 2024-11-19 21:05 | 0.63552 | 0 | 0.5286 | 0.0033994 | 1 | #N/A | 0.52522 | 0.5252 | 0.52519 | 0.52518 | 0.54219
| 4875000 | 1001 | 2024-11-19 21:06 | 0.62504 | 0 | 0.53882 | 0.0033989 |  | 0.63713 | 0.54221 | 0.54222 | 0.54222 | 0.54223 | 0.52523
| 1000 | 1000 | 2024-11-19 21:06 | 0.63552 | 0 | 0.5286 | 0.0033998 | 1 | #N/A | 0.52521 | 0.5252 | 0.52519 | 0.52518 | 0.54219
| 4905000 | 981 | 2024-11-19 21:07 | 0.62375 | 0 | 0.53879 | 0.003437 |  | 0.63595 | 0.5422 | 0.54222 | 0.54223 | 0.54224 | 0.52504
| 1000 | 1000 | 2024-11-19 21:08 | 0.63552 | 0 | 0.52845 | 0.0034379 | 1 | #N/A | 0.52503 | 0.52501 | 0.525 | 0.525 | 0.5422
| 4915000 | 983 | 2024-11-19 21:09 | 0.62335 | 0 | 0.53878 | 0.0034428 |  | 0.63555 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.52501
| 1000 | 1000 | 2024-11-19 21:09 | 0.63554 | 0 | 0.52844 | 0.0034451 | 1 | #N/A | 0.525 | 0.525 | 0.52499 | 0.52498 | 0.54222
| 4925000 | 985 | 2024-11-19 21:10 | 0.6229 | 0 | 0.53879 | 0.0034553 |  | 0.63515 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.52497
| 1000 | 1000 | 2024-11-19 21:11 | 0.63553 | 0 | 0.5284 | 0.0034518 | 1 | #N/A | 0.52496 | 0.52495 | 0.52494 | 0.52493 | 0.5422
| 4935000 | 987 | 2024-11-19 21:12 | 0.62245 | 0 | 0.53876 | 0.0034647 |  | 0.63472 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.5249
| 1000 | 1000 | 2024-11-19 21:12 | 0.63554 | 0 | 0.52834 | 0.0034677 | 1 | #N/A | 0.52489 | 0.52488 | 0.52487 | 0.52486 | 0.54222
| 4940000 | 988 | 2024-11-19 21:13 | 0.62222 | 0 | 0.53877 | 0.0034695 |  | 0.6345 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.52489
| 1000 | 1000 | 2024-11-19 21:14 | 0.63555 | 0 | 0.52834 | 0.0034709 | 1 | #N/A | 0.52489 | 0.52488 | 0.52486 | 0.52485 | 0.54222
| 4950000 | 990 | 2024-11-19 21:14 | 0.62181 | 0 | 0.53877 | 0.0034795 |  | 0.63412 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.52486
| 1000 | 1000 | 2024-11-19 21:15 | 0.63557 | 0 | 0.52831 | 0.0034813 | 1 | #N/A | 0.52485 | 0.52484 | 0.52483 | 0.52482 | 0.54224
| 4960000 | 992 | 2024-11-19 21:16 | 0.62128 | 0 | 0.53876 | 0.0034986 |  | 0.63367 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.52476
| 1000 | 1000 | 2024-11-19 21:17 | 0.63554 | 0 | 0.52823 | 0.0034966 | 1 | #N/A | 0.52475 | 0.52474 | 0.52473 | 0.52472 | 0.54222
| 4970000 | 994 | 2024-11-19 21:17 | 0.62073 | 0 | 0.53872 | 0.0035194 |  | 0.63318 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.52464
| 1000 | 1000 | 2024-11-19 21:18 | 0.63552 | 0 | 0.52813 | 0.0035159 | 1 | #N/A | 0.52463 | 0.52462 | 0.5246 | 0.52459 | 0.54219
| 4980000 | 996 | 2024-11-19 21:19 | 0.6203 | 0 | 0.5387 | 0.0035202 |  | 0.63273 | 0.5422 | 0.54221 | 0.54223 | 0.54224 | 0.52462
| 1000 | 1000 | 2024-11-19 21:20 | 0.63554 | 0 | 0.52812 | 0.0035244 | 1 | #N/A | 0.5246 | 0.52459 | 0.52459 | 0.52458 | 0.54221
| 4990000 | 998 | 2024-11-19 21:20 | 0.61977 | 0 | 0.5387 | 0.0035458 |  | 0.6323 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.52452
| 1000 | 1000 | 2024-11-19 21:21 | 0.63553 | 0 | 0.52803 | 0.0035443 | 1 | #N/A | 0.5245 | 0.52449 | 0.52448 | 0.52447 | 0.54221
| 5000000 | 1000 | 2024-11-19 21:22 | 0.61931 | 0 | 0.53868 | 0.0035492 |  | 0.63183 | 0.54221 | 0.54223 | 0.54223 | 0.54225 | 0.52448
| 1000 | 1000 | 2024-11-19 21:23 | 0.63553 | 0 | 0.52801 | 0.0035495 | 1 | #N/A | 0.52447 | 0.52447 | 0.52445 | 0.52445 | 0.54221
| 5010000 | 1002 | 2024-11-19 21:23 | 0.61881 | 0 | 0.53867 | 0.0035648 |  | 0.63139 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.52441
| 1000 | 1000 | 2024-11-19 21:24 | 0.63552 | 0 | 0.52794 | 0.003565 | 1 | #N/A | 0.52439 | 0.52438 | 0.52437 | 0.52436 | 0.5422
| 5020000 | 1004 | 2024-11-19 21:25 | 0.61839 | 0 | 0.53866 | 0.0035727 |  | 0.63098 | 0.54222 | 0.54224 | 0.54224 | 0.54225 | 0.52437
| 1000 | 1000 | 2024-11-19 21:26 | 0.63553 | 0 | 0.52792 | 0.0035716 | 1 | #N/A | 0.52436 | 0.52435 | 0.52435 | 0.52434 | 0.54221
| 5030000 | 1006 | 2024-11-19 21:26 | 0.61788 | 0 | 0.53864 | 0.0035873 |  | 0.63051 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.52429
| 1000 | 1000 | 2024-11-19 21:27 | 0.63553 | 0 | 0.52785 | 0.003587 | 1 | #N/A | 0.52428 | 0.52427 | 0.52426 | 0.52425 | 0.5422
| 5040000 | 1008 | 2024-11-19 21:28 | 0.61742 | 0 | 0.53864 | 0.003595 |  | 0.63007 | 0.54221 | 0.54222 | 0.54224 | 0.54224 | 0.52426
| 1000 | 1000 | 2024-11-19 21:29 | 0.63553 | 0 | 0.52782 | 0.003596 | 1 | #N/A | 0.52424 | 0.52423 | 0.52422 | 0.52421 | 0.5422
| 5050000 | 1010 | 2024-11-19 21:29 | 0.61695 | 0 | 0.53863 | 0.0036109 |  | 0.62965 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.52418
| 1000 | 1000 | 2024-11-19 21:30 | 0.63553 | 0 | 0.52776 | 0.0036099 | 1 | #N/A | 0.52417 | 0.52416 | 0.52415 | 0.52414 | 0.5422
| 5060000 | 1012 | 2024-11-19 21:31 | 0.61653 | 0 | 0.53861 | 0.0036161 |  | 0.62922 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.52414
| 1000 | 1000 | 2024-11-19 21:32 | 0.63556 | 0 | 0.52774 | 0.0036205 | 1 | #N/A | 0.52413 | 0.52413 | 0.52412 | 0.52411 | 0.54223
| 5065000 | 1013 | 2024-11-19 21:32 | 0.61623 | 0 | 0.53861 | 0.0036343 |  | 0.62901 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.52407
| 1000 | 1000 | 2024-11-19 21:33 | 0.6355 | 0 | 0.52766 | 0.0036301 | 1 | #N/A | 0.52405 | 0.52404 | 0.52402 | 0.52401 | 0.54218
| 5075000 | 1015 | 2024-11-19 21:34 | 0.6158 | 0 | 0.53857 | 0.0036382 |  | 0.62856 | 0.54219 | 0.5422 | 0.54221 | 0.54222 | 0.52401
| 1000 | 1000 | 2024-11-19 21:34 | 0.63553 | 0 | 0.52764 | 0.0036401 | 1 | #N/A | 0.52401 | 0.524 | 0.524 | 0.524 | 0.5422
| 5100000 | 1003 | 2024-11-19 21:35 | 0.61478 | 0 | 0.53857 | 0.0036637 |  | 0.62762 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.52392
| 1000 | 1000 | 2024-11-19 21:36 | 0.63554 | 0 | 0.52755 | 0.0036645 | 1 | #N/A | 0.5239 | 0.52389 | 0.52388 | 0.52387 | 0.54221
| 5100000 | 1003 | 2024-11-19 21:37 | 0.6148 | 0 | 0.53857 | 0.003663 |  | 0.62764 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.52392
| 1000 | 1000 | 2024-11-19 21:37 | 0.63554 | 0 | 0.52756 | 0.0036638 | 1 | #N/A | 0.52391 | 0.5239 | 0.52389 | 0.52388 | 0.54222
| 5100000 | 1003 | 2024-11-19 21:38 | 0.61479 | 0 | 0.53858 | 0.0036622 |  | 0.62764 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.52393
| 1000 | 1000 | 2024-11-19 21:39 | 0.63555 | 0 | 0.52756 | 0.0036665 | 1 | #N/A | 0.52391 | 0.5239 | 0.52388 | 0.52387 | 0.54222
| 5125000 | 984 | 2024-11-19 21:40 | 0.6138 | 0 | 0.53855 | 0.0036951 |  | 0.62672 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.52377
| 1000 | 1000 | 2024-11-19 21:40 | 0.63555 | 0 | 0.52744 | 0.0036954 | 1 | #N/A | 0.52376 | 0.52375 | 0.52374 | 0.52373 | 0.54222
| 5125000 | 984 | 2024-11-19 21:41 | 0.61376 | 0 | 0.53855 | 0.0036981 |  | 0.62671 | 0.54224 | 0.54224 | 0.54225 | 0.54225 | 0.52375
| 1000 | 1000 | 2024-11-19 21:42 | 0.63553 | 0 | 0.52743 | 0.0036945 | 1 | #N/A | 0.52374 | 0.52373 | 0.52373 | 0.52372 | 0.5422
| 5125000 | 984 | 2024-11-19 21:43 | 0.61378 | 0 | 0.53854 | 0.0036948 |  | 0.62671 | 0.54221 | 0.54223 | 0.54224 | 0.54225 | 0.52376
| 1000 | 1000 | 2024-11-19 21:43 | 0.63554 | 0 | 0.52742 | 0.0036972 | 1 | #N/A | 0.52374 | 0.52373 | 0.52372 | 0.52371 | 0.54221
| 5160000 | 989 | 2024-11-19 21:44 | 0.61255 | 0 | 0.53849 | 0.0037376 |  | 0.62561 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.52354
| 1000 | 1000 | 2024-11-19 21:45 | 0.63553 | 0 | 0.52725 | 0.0037383 | 1 | #N/A | 0.52353 | 0.52352 | 0.52351 | 0.5235 | 0.54221
| 5160000 | 989 | 2024-11-19 21:46 | 0.61254 | 0 | 0.53851 | 0.0037395 |  | 0.6256 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.52355
| 1000 | 1000 | 2024-11-19 21:46 | 0.63554 | 0 | 0.52726 | 0.0037374 | 1 | #N/A | 0.52354 | 0.52353 | 0.52352 | 0.52351 | 0.54221
| 5160000 | 989 | 2024-11-19 21:47 | 0.61255 | 0 | 0.5385 | 0.0037361 |  | 0.6256 | 0.54222 | 0.54222 | 0.54224 | 0.54225 | 0.52355
| 1000 | 1000 | 2024-11-19 21:48 | 0.63553 | 0 | 0.52726 | 0.0037362 | 1 | #N/A | 0.52354 | 0.52353 | 0.52352 | 0.52351 | 0.54221
| 5160000 | 989 | 2024-11-19 21:48 | 0.61256 | 0 | 0.5385 | 0.003734 |  | 0.6256 | 0.54221 | 0.54223 | 0.54224 | 0.54225 | 0.52356
| 1000 | 1000 | 2024-11-19 21:49 | 0.63555 | 0 | 0.52728 | 0.0037375 | 1 | #N/A | 0.52355 | 0.52355 | 0.52353 | 0.52353 | 0.54223
| 5200000 | 988 | 2024-11-19 21:50 | 0.61141 | 0 | 0.53846 | 0.0037846 |  | 0.62462 | 0.54223 | 0.54224 | 0.54226 | 0.54227 | 0.52333
| 1000 | 1000 | 2024-11-19 21:51 | 0.63554 | 0 | 0.52708 | 0.0037838 | 1 | #N/A | 0.52332 | 0.5233 | 0.52329 | 0.52328 | 0.54221
| 5200000 | 988 | 2024-11-19 21:51 | 0.61137 | 0 | 0.53845 | 0.0037861 |  | 0.62458 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.5233
| 1000 | 1000 | 2024-11-19 21:52 | 0.63554 | 0 | 0.52706 | 0.003787 | 1 | #N/A | 0.52329 | 0.52328 | 0.52327 | 0.52326 | 0.54221
| 5200000 | 988 | 2024-11-19 21:53 | 0.6114 | 0 | 0.53845 | 0.0037853 |  | 0.6246 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.52331
| 1000 | 1000 | 2024-11-19 21:54 | 0.63556 | 0 | 0.52707 | 0.0037891 | 1 | #N/A | 0.5233 | 0.52328 | 0.52327 | 0.52327 | 0.54223
| 5200000 | 988 | 2024-11-19 21:54 | 0.61138 | 0 | 0.53846 | 0.0037887 |  | 0.6246 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.5233
| 1000 | 1000 | 2024-11-19 21:55 | 0.63553 | 0 | 0.52706 | 0.0037854 | 1 | #N/A | 0.5233 | 0.52328 | 0.52327 | 0.52326 | 0.5422
| 5250000 | 1001 | 2024-11-19 21:56 | 0.6104 | 0 | 0.53839 | 0.0038478 |  | 0.62382 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.523
| 1000 | 1000 | 2024-11-19 21:57 | 0.63554 | 0 | 0.52681 | 0.0038503 | 1 | #N/A | 0.52298 | 0.52296 | 0.52295 | 0.52294 | 0.54221
| 5250000 | 1001 | 2024-11-19 21:57 | 0.61038 | 0 | 0.53838 | 0.0038491 |  | 0.6238 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.52299
| 1000 | 1000 | 2024-11-19 21:58 | 0.63552 | 0 | 0.5268 | 0.003848 | 1 | #N/A | 0.52298 | 0.52296 | 0.52295 | 0.52294 | 0.5422
| 5250000 | 1001 | 2024-11-19 21:59 | 0.61038 | 0 | 0.53837 | 0.003854 |  | 0.62381 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.52296
| 1000 | 1000 | 2024-11-19 22:00 | 0.63552 | 0 | 0.52679 | 0.0038523 | 1 | #N/A | 0.52295 | 0.52294 | 0.52293 | 0.52292 | 0.54219
| 5250000 | 1001 | 2024-11-19 22:00 | 0.6104 | 0 | 0.53838 | 0.0038487 |  | 0.62381 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.52298
| 1000 | 1000 | 2024-11-19 22:01 | 0.63551 | 0 | 0.5268 | 0.0038478 | 1 | #N/A | 0.52297 | 0.52296 | 0.52295 | 0.52294 | 0.5422
| 5280000 | 1012 | 2024-11-19 22:02 | 0.6101 | 0 | 0.53834 | 0.003882 |  | 0.62362 | 0.54221 | 0.54222 | 0.54222 | 0.54224 | 0.52281
| 1000 | 1000 | 2024-11-19 22:03 | 0.63552 | 0 | 0.52667 | 0.0038818 | 1 | #N/A | 0.5228 | 0.52279 | 0.52278 | 0.52278 | 0.5422
| 5280000 | 1012 | 2024-11-19 22:03 | 0.61008 | 0 | 0.53834 | 0.0038814 |  | 0.6236 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.52282
| 1000 | 1000 | 2024-11-19 22:04 | 0.63553 | 0 | 0.52667 | 0.0038832 | 1 | #N/A | 0.5228 | 0.5228 | 0.52278 | 0.52278 | 0.54221
| 5300000 | 1007 | 2024-11-19 22:05 | 0.61006 | 0 | 0.53834 | 0.0039011 |  | 0.62364 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.52273
| 1000 | 1000 | 2024-11-19 22:06 | 0.63555 | 0 | 0.52661 | 0.0039014 | 1 | #N/A | 0.52272 | 0.52272 | 0.52271 | 0.5227 | 0.54222
| 5300000 | 1007 | 2024-11-19 22:06 | 0.61007 | 0 | 0.53834 | 0.0039028 |  | 0.62365 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.52273
| 1000 | 1000 | 2024-11-19 22:07 | 0.63556 | 0 | 0.52661 | 0.0039043 | 1 | #N/A | 0.52272 | 0.52271 | 0.52271 | 0.5227 | 0.54223
| 5300000 | 1007 | 2024-11-19 22:08 | 0.61008 | 0 | 0.53835 | 0.0039005 |  | 0.62366 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.52275
| 1000 | 1000 | 2024-11-19 22:08 | 0.63556 | 0 | 0.52663 | 0.0039017 | 1 | #N/A | 0.52273 | 0.52273 | 0.52272 | 0.52271 | 0.54223
| 5325000 | 994 | 2024-11-19 22:09 | 0.61008 | 0 | 0.53832 | 0.0039405 |  | 0.62381 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.52255
| 1000 | 1000 | 2024-11-19 22:10 | 0.63554 | 0 | 0.52647 | 0.0039372 | 1 | #N/A | 0.52255 | 0.52253 | 0.52253 | 0.52252 | 0.54222
| 5325000 | 994 | 2024-11-19 22:11 | 0.61005 | 0 | 0.5383 | 0.0039397 |  | 0.62378 | 0.54223 | 0.54223 | 0.54224 | 0.54225 | 0.52254
| 1000 | 1000 | 2024-11-19 22:11 | 0.63554 | 0 | 0.52645 | 0.0039398 | 1 | #N/A | 0.52253 | 0.52252 | 0.52251 | 0.52249 | 0.54221
| 5325000 | 994 | 2024-11-19 22:12 | 0.61004 | 0 | 0.5383 | 0.003941 |  | 0.62378 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.52253
| 1000 | 1000 | 2024-11-19 22:13 | 0.63552 | 0 | 0.52645 | 0.0039385 | 1 | #N/A | 0.52252 | 0.52251 | 0.52251 | 0.52249 | 0.5422
| 5340000 | 979 | 2024-11-19 22:14 | 0.61015 | 0 | 0.53827 | 0.0039564 |  | 0.62396 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.52244
| 1000 | 1000 | 2024-11-19 22:14 | 0.63551 | 0 | 0.52636 | 0.0039571 | 1 | #N/A | 0.52243 | 0.52241 | 0.5224 | 0.52239 | 0.54219
| 5375000 | 989 | 2024-11-19 22:15 | 0.61057 | 0 | 0.53822 | 0.0040031 |  | 0.62452 | 0.5422 | 0.54222 | 0.54223 | 0.54224 | 0.52221
| 1000 | 1000 | 2024-11-19 22:16 | 0.63553 | 0 | 0.52619 | 0.0040027 | 1 | #N/A | 0.5222 | 0.52219 | 0.52219 | 0.52218 | 0.5422
| 5375000 | 989 | 2024-11-19 22:17 | 0.61053 | 0 | 0.53822 | 0.004008 |  | 0.6245 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.52218
| 1000 | 1000 | 2024-11-19 22:17 | 0.63553 | 0 | 0.52617 | 0.0040077 | 1 | #N/A | 0.52217 | 0.52216 | 0.52216 | 0.52216 | 0.5422
| 5375000 | 989 | 2024-11-19 22:18 | 0.61056 | 0 | 0.53822 | 0.0040032 |  | 0.62452 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.52221
| 1000 | 1000 | 2024-11-19 22:19 | 0.63553 | 0 | 0.52619 | 0.0040031 | 1 | #N/A | 0.5222 | 0.5222 | 0.52218 | 0.52217 | 0.5422
| 5400000 | 999 | 2024-11-19 22:20 | 0.61108 | 0 | 0.5382 | 0.0040298 |  | 0.62515 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.52208
| 1000 | 1000 | 2024-11-19 22:20 | 0.63553 | 0 | 0.52608 | 0.0040303 | 1 | #N/A | 0.52207 | 0.52206 | 0.52205 | 0.52203 | 0.5422
| 5400000 | 999 | 2024-11-19 22:21 | 0.61108 | 0 | 0.53819 | 0.0040305 |  | 0.62515 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.52207
| 1000 | 1000 | 2024-11-19 22:22 | 0.63553 | 0 | 0.52608 | 0.0040319 | 1 | #N/A | 0.52206 | 0.52205 | 0.52204 | 0.52203 | 0.54221
| 5400000 | 999 | 2024-11-19 22:22 | 0.61105 | 0 | 0.5382 | 0.0040315 |  | 0.62513 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.52208
| 1000 | 1000 | 2024-11-19 22:23 | 0.63553 | 0 | 0.52608 | 0.0040315 | 1 | #N/A | 0.52207 | 0.52205 | 0.52204 | 0.52204 | 0.54221
| 5400000 | 999 | 2024-11-19 22:24 | 0.61107 | 0 | 0.5382 | 0.0040332 |  | 0.62516 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.52207
| 1000 | 1000 | 2024-11-19 22:25 | 0.63552 | 0 | 0.52607 | 0.004031 | 1 | #N/A | 0.52205 | 0.52205 | 0.52204 | 0.52203 | 0.5422
| 5450000 | 981 | 2024-11-19 22:25 | 0.61245 | 0 | 0.53814 | 0.0040927 |  | 0.6268 | 0.54221 | 0.54223 | 0.54224 | 0.54224 | 0.52177
| 1000 | 1000 | 2024-11-19 22:26 | 0.63553 | 0 | 0.52582 | 0.004096 | 1 | #N/A | 0.52175 | 0.52173 | 0.52172 | 0.5217 | 0.5422
| 5460000 | 1001 | 2024-11-19 22:27 | 0.61276 | 0 | 0.53812 | 0.0041055 |  | 0.62715 | 0.54222 | 0.54222 | 0.54223 | 0.54224 | 0.5217
| 1000 | 1000 | 2024-11-19 22:28 | 0.63553 | 0 | 0.52578 | 0.0041052 | 1 | #N/A | 0.52169 | 0.52168 | 0.52167 | 0.52166 | 0.5422
| 5460000 | 1001 | 2024-11-19 22:28 | 0.61276 | 0 | 0.53813 | 0.0041054 |  | 0.62714 | 0.54221 | 0.54223 | 0.54224 | 0.54224 | 0.52171
| 1000 | 1000 | 2024-11-19 22:29 | 0.63554 | 0 | 0.52579 | 0.004106 | 1 | #N/A | 0.5217 | 0.52168 | 0.52167 | 0.52167 | 0.54221
| 5460000 | 1001 | 2024-11-19 22:30 | 0.61277 | 0 | 0.53814 | 0.0041068 |  | 0.62715 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.52171
| 1000 | 1000 | 2024-11-19 22:31 | 0.63555 | 0 | 0.5258 | 0.0041075 | 1 | #N/A | 0.5217 | 0.5217 | 0.52168 | 0.52167 | 0.54223
| 5460000 | 1001 | 2024-11-19 22:31 | 0.61277 | 0 | 0.53814 | 0.0041103 |  | 0.62718 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.5217
| 1000 | 1000 | 2024-11-19 22:32 | 0.63554 | 0 | 0.52579 | 0.0041072 | 1 | #N/A | 0.52169 | 0.52168 | 0.52168 | 0.52167 | 0.54222
| 5460000 | 1001 | 2024-11-19 22:33 | 0.61275 | 0 | 0.53814 | 0.0041087 |  | 0.62715 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.5217
| 1000 | 1000 | 2024-11-19 22:34 | 0.63553 | 0 | 0.52579 | 0.0041051 | 1 | #N/A | 0.52169 | 0.52168 | 0.52168 | 0.52167 | 0.54221
| 5500000 | 1001 | 2024-11-19 22:34 | 0.6143 | 0 | 0.53809 | 0.0041451 |  | 0.62886 | 0.54221 | 0.54223 | 0.54224 | 0.54225 | 0.52151
| 1000 | 1000 | 2024-11-19 22:35 | 0.63557 | 0 | 0.52563 | 0.0041508 | 1 | #N/A | 0.52149 | 0.52148 | 0.52148 | 0.52147 | 0.54223
| 5500000 | 1001 | 2024-11-19 22:36 | 0.61429 | 0 | 0.53811 | 0.0041513 |  | 0.62887 | 0.54224 | 0.54226 | 0.54226 | 0.54227 | 0.5215
| 1000 | 1000 | 2024-11-19 22:37 | 0.63557 | 0 | 0.52563 | 0.0041497 | 1 | #N/A | 0.52149 | 0.52149 | 0.52148 | 0.52147 | 0.54223
| 5500000 | 1001 | 2024-11-19 22:37 | 0.61429 | 0 | 0.53811 | 0.004151 |  | 0.62887 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.5215
| 1000 | 1000 | 2024-11-19 22:38 | 0.63556 | 0 | 0.52563 | 0.0041493 | 1 | #N/A | 0.5215 | 0.52149 | 0.52147 | 0.52147 | 0.54223
| 5520000 | 1012 | 2024-11-19 22:39 | 0.61507 | 0 | 0.53807 | 0.0041788 |  | 0.62977 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.52136
| 1000 | 1000 | 2024-11-19 22:40 | 0.63555 | 0 | 0.52552 | 0.0041764 | 1 | #N/A | 0.52135 | 0.52134 | 0.52134 | 0.52133 | 0.54222
| 5520000 | 1012 | 2024-11-19 22:40 | 0.61507 | 0 | 0.53806 | 0.0041791 |  | 0.62977 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.52134
| 1000 | 1000 | 2024-11-19 22:41 | 0.63556 | 0 | 0.5255 | 0.0041807 | 1 | #N/A | 0.52134 | 0.52133 | 0.52132 | 0.5213 | 0.54222
| 5550000 | 999 | 2024-11-19 22:42 | 0.61637 | 0 | 0.53804 | 0.0042143 |  | 0.63124 | 0.54224 | 0.54225 | 0.54226 | 0.54228 | 0.52118
| 1000 | 1000 | 2024-11-19 22:42 | 0.63555 | 0 | 0.52536 | 0.0042131 | 1 | #N/A | 0.52117 | 0.52115 | 0.52115 | 0.52114 | 0.54222
| 5550000 | 999 | 2024-11-19 22:43 | 0.61637 | 0 | 0.53803 | 0.0042162 |  | 0.63124 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.52116
| 1000 | 1000 | 2024-11-19 22:44 | 0.63553 | 0 | 0.52535 | 0.004214 | 1 | #N/A | 0.52115 | 0.52114 | 0.52113 | 0.52112 | 0.5422
| 5550000 | 999 | 2024-11-19 22:45 | 0.61637 | 0 | 0.53802 | 0.0042141 |  | 0.63124 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.52116
| 1000 | 1000 | 2024-11-19 22:45 | 0.63554 | 0 | 0.52535 | 0.0042176 | 1 | #N/A | 0.52115 | 0.52113 | 0.52112 | 0.52112 | 0.54222
| 5550000 | 999 | 2024-11-19 22:46 | 0.61635 | 0 | 0.53803 | 0.0042195 |  | 0.63123 | 0.54224 | 0.54225 | 0.54226 | 0.54226 | 0.52115
| 1000 | 1000 | 2024-11-19 22:47 | 0.63553 | 0 | 0.52535 | 0.0042159 | 1 | #N/A | 0.52114 | 0.52113 | 0.52113 | 0.52112 | 0.54221
| 5600000 | 1008 | 2024-11-19 22:48 | 0.61854 | 0 | 0.53796 | 0.0042699 |  | 0.63366 | 0.54222 | 0.54223 | 0.54223 | 0.54225 | 0.52088
| 1000 | 1000 | 2024-11-19 22:48 | 0.63555 | 0 | 0.52513 | 0.0042726 | 1 | #N/A | 0.52087 | 0.52086 | 0.52085 | 0.52084 | 0.54222
| 5600000 | 1008 | 2024-11-19 22:49 | 0.61858 | 0 | 0.53797 | 0.0042672 |  | 0.63368 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.5209
| 1000 | 1000 | 2024-11-19 22:50 | 0.63556 | 0 | 0.52515 | 0.0042697 | 1 | #N/A | 0.52089 | 0.52088 | 0.52087 | 0.52086 | 0.54223
| 5600000 | 1008 | 2024-11-19 22:51 | 0.61855 | 0 | 0.53798 | 0.0042729 |  | 0.63367 | 0.54224 | 0.54225 | 0.54226 | 0.54227 | 0.52089
| 1000 | 1000 | 2024-11-19 22:51 | 0.63558 | 0 | 0.52514 | 0.0042747 | 1 | #N/A | 0.52088 | 0.52087 | 0.52087 | 0.52086 | 0.54224
| 5600000 | 1008 | 2024-11-19 22:52 | 0.61856 | 0 | 0.53799 | 0.0042729 |  | 0.63368 | 0.54225 | 0.54225 | 0.54226 | 0.54227 | 0.52089
| 1000 | 1000 | 2024-11-19 22:53 | 0.63557 | 0 | 0.52515 | 0.004274 | 1 | #N/A | 0.52088 | 0.52088 | 0.52087 | 0.52087 | 0.54224
| 5625000 | 999 | 2024-11-19 22:54 | 0.61959 | 0 | 0.53796 | 0.0043099 |  | 0.63488 | 0.54225 | 0.54226 | 0.54227 | 0.54228 | 0.52072
| 1000 | 1000 | 2024-11-19 22:54 | 0.63555 | 0 | 0.525 | 0.0043058 | 1 | #N/A | 0.52071 | 0.5207 | 0.52069 | 0.52068 | 0.54222
| 5625000 | 999 | 2024-11-19 22:55 | 0.61957 | 0 | 0.53794 | 0.0043091 |  | 0.63485 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.5207
| 1000 | 1000 | 2024-11-19 22:56 | 0.63555 | 0 | 0.52498 | 0.00431 | 1 | #N/A | 0.52069 | 0.52068 | 0.52067 | 0.52066 | 0.54222
| 5640000 | 987 | 2024-11-19 22:56 | 0.62018 | 0 | 0.53792 | 0.0043261 |  | 0.63556 | 0.54223 | 0.54224 | 0.54225 | 0.54227 | 0.52062
| 1000 | 1000 | 2024-11-19 22:57 | 0.63553 | 0 | 0.52491 | 0.0043266 | 1 | #N/A | 0.5206 | 0.52059 | 0.52057 | 0.52056 | 0.54221
| 5640000 | 987 | 2024-11-19 22:58 | 0.62018 | 0 | 0.53791 | 0.0043266 |  | 0.63555 | 0.54222 | 0.54223 | 0.54224 | 0.54226 | 0.5206
| 1000 | 1000 | 2024-11-19 22:59 | 0.63553 | 0 | 0.5249 | 0.0043257 | 1 | #N/A | 0.52059 | 0.52058 | 0.52057 | 0.52056 | 0.5422
| 5650000 | 1017 | 2024-11-19 23:00 | 0.62054 | 0 | 0.53789 | 0.0043349 |  | 0.63594 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.52055
| 1000 | 1000 | 2024-11-19 23:00 | 0.63554 | 0 | 0.52486 | 0.0043371 | 1 | #N/A | 0.52053 | 0.52053 | 0.52052 | 0.52052 | 0.54221
| 5680000 | 994 | 2024-11-19 23:01 | 0.62157 | 0 | 0.53787 | 0.004376 |  | 0.63715 | 0.54222 | 0.54223 | 0.54225 | 0.54226 | 0.52036
| 1000 | 1000 | 2024-11-19 23:02 | 0.63554 | 0 | 0.52472 | 0.0043748 | 1 | #N/A | 0.52035 | 0.52035 | 0.52034 | 0.52033 | 0.54222
| 5680000 | 994 | 2024-11-19 23:03 | 0.62155 | 0 | 0.53786 | 0.0043785 |  | 0.63715 | 0.54222 | 0.54224 | 0.54225 | 0.54226 | 0.52035
| 1000 | 1000 | 2024-11-19 23:03 | 0.63553 | 0 | 0.5247 | 0.0043777 | 1 | #N/A | 0.52034 | 0.52033 | 0.52031 | 0.5203 | 0.54221
| 5680000 | 994 | 2024-11-19 23:04 | 0.62154 | 0 | 0.53786 | 0.004378 |  | 0.63714 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.52034
| 1000 | 1000 | 2024-11-19 23:05 | 0.63554 | 0 | 0.52469 | 0.0043808 | 1 | #N/A | 0.52033 | 0.52031 | 0.52031 | 0.5203 | 0.54221
| 5700000 | 1007 | 2024-11-19 23:06 | 0.62209 | 0 | 0.53783 | 0.0044044 |  | 0.6378 | 0.54223 | 0.54223 | 0.54224 | 0.54225 | 0.52022
| 1000 | 1000 | 2024-11-19 23:06 | 0.63554 | 0 | 0.52459 | 0.0044053 | 1 | #N/A | 0.5202 | 0.52019 | 0.52018 | 0.52016 | 0.54221
| 5720000 | 1001 | 2024-11-19 23:07 | 0.62253 | 0 | 0.53781 | 0.0044311 |  | 0.63835 | 0.54222 | 0.54224 | 0.54224 | 0.54225 | 0.52008
| 1000 | 1000 | 2024-11-19 23:08 | 0.63551 | 0 | 0.52448 | 0.0044265 | 1 | #N/A | 0.52007 | 0.52006 | 0.52005 | 0.52004 | 0.54219
| 5720000 | 1001 | 2024-11-19 23:09 | 0.62252 | 0 | 0.53779 | 0.0044294 |  | 0.63831 | 0.5422 | 0.54222 | 0.54223 | 0.54224 | 0.52007
| 1000 | 1000 | 2024-11-19 23:09 | 0.63554 | 0 | 0.52448 | 0.0044338 | 1 | #N/A | 0.52006 | 0.52005 | 0.52004 | 0.52003 | 0.54222
| 5720000 | 1001 | 2024-11-19 23:10 | 0.62253 | 0 | 0.53781 | 0.0044311 |  | 0.63834 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.52009
| 1000 | 1000 | 2024-11-19 23:11 | 0.63552 | 0 | 0.52449 | 0.004427 | 1 | #N/A | 0.52008 | 0.52007 | 0.52005 | 0.52004 | 0.54219
| 5750000 | 989 | 2024-11-19 23:12 | 0.62292 | 0 | 0.53776 | 0.0044659 |  | 0.63887 | 0.54221 | 0.54222 | 0.54222 | 0.54224 | 0.51989
| 1000 | 1000 | 2024-11-19 23:12 | 0.63553 | 0 | 0.52433 | 0.0044695 | 1 | #N/A | 0.51988 | 0.51986 | 0.51985 | 0.51985 | 0.54221
| 5750000 | 989 | 2024-11-19 23:13 | 0.62292 | 0 | 0.53777 | 0.0044664 |  | 0.63888 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.5199
| 1000 | 1000 | 2024-11-19 23:14 | 0.63551 | 0 | 0.52433 | 0.0044645 | 1 | #N/A | 0.51989 | 0.51987 | 0.51986 | 0.51985 | 0.54219
| 5760000 | 1008 | 2024-11-19 23:15 | 0.623 | 0 | 0.53775 | 0.0044695 |  | 0.63895 | 0.5422 | 0.54221 | 0.54222 | 0.54224 | 0.51987
| 1000 | 1000 | 2024-11-19 23:15 | 0.63553 | 0 | 0.52432 | 0.0044716 | 1 | #N/A | 0.51986 | 0.51985 | 0.51985 | 0.51984 | 0.54221
| 5775000 | 1001 | 2024-11-19 23:16 | 0.62296 | 0 | 0.53774 | 0.0044991 |  | 0.63904 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.51974
| 1000 | 1000 | 2024-11-19 23:17 | 0.63553 | 0 | 0.52421 | 0.0044984 | 1 | #N/A | 0.51973 | 0.51972 | 0.51971 | 0.51971 | 0.54221
| 5775000 | 1001 | 2024-11-19 23:18 | 0.62295 | 0 | 0.53773 | 0.0044997 |  | 0.63903 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.51973
| 1000 | 1000 | 2024-11-19 23:18 | 0.63551 | 0 | 0.5242 | 0.0044978 | 1 | #N/A | 0.51972 | 0.51971 | 0.51969 | 0.51968 | 0.54219
| 5775000 | 1001 | 2024-11-19 23:19 | 0.62295 | 0 | 0.53773 | 0.0044973 |  | 0.63902 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.51974
| 1000 | 1000 | 2024-11-19 23:20 | 0.63551 | 0 | 0.52421 | 0.004496 | 1 | #N/A | 0.51973 | 0.51972 | 0.51971 | 0.5197 | 0.54219
| 5800000 | 986 | 2024-11-19 23:21 | 0.62274 | 0 | 0.53769 | 0.0045226 |  | 0.63888 | 0.5422 | 0.54221 | 0.54222 | 0.54223 | 0.5196
| 1000 | 1000 | 2024-11-19 23:21 | 0.63554 | 0 | 0.5241 | 0.0045267 | 1 | #N/A | 0.51959 | 0.51958 | 0.51957 | 0.51956 | 0.54221
| 5800000 | 986 | 2024-11-19 23:22 | 0.62271 | 0 | 0.53771 | 0.0045291 |  | 0.63889 | 0.54222 | 0.54224 | 0.54225 | 0.54225 | 0.51959
| 1000 | 1000 | 2024-11-19 23:23 | 0.63553 | 0 | 0.52409 | 0.0045281 | 1 | #N/A | 0.51958 | 0.51957 | 0.51956 | 0.51955 | 0.54221
| 5800000 | 986 | 2024-11-19 23:24 | 0.62271 | 0 | 0.53772 | 0.0045304 |  | 0.63888 | 0.54222 | 0.54224 | 0.54226 | 0.54227 | 0.51959
| 1000 | 1000 | 2024-11-19 23:24 | 0.63555 | 0 | 0.5241 | 0.0045291 | 1 | #N/A | 0.51958 | 0.51958 | 0.51957 | 0.51956 | 0.54222
| 5840000 | 1022 | 2024-11-19 23:25 | 0.62161 | 0 | 0.53767 | 0.0045699 |  | 0.63791 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.51939
| 1000 | 1000 | 2024-11-19 23:26 | 0.63554 | 0 | 0.52393 | 0.0045704 | 1 | #N/A | 0.51938 | 0.51937 | 0.51936 | 0.51935 | 0.54222
| 5850000 | 1014 | 2024-11-19 23:27 | 0.62118 | 0 | 0.53765 | 0.004582 |  | 0.63751 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.51933
| 1000 | 1000 | 2024-11-19 23:27 | 0.63554 | 0 | 0.52389 | 0.0045821 | 1 | #N/A | 0.51932 | 0.51931 | 0.5193 | 0.51929 | 0.54222
| 5850000 | 1014 | 2024-11-19 23:28 | 0.6212 | 0 | 0.53766 | 0.0045824 |  | 0.63753 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.51933
| 1000 | 1000 | 2024-11-19 23:29 | 0.63554 | 0 | 0.52389 | 0.0045823 | 1 | #N/A | 0.51932 | 0.51931 | 0.5193 | 0.51929 | 0.54221
| 5850000 | 1014 | 2024-11-19 23:30 | 0.62121 | 0 | 0.53766 | 0.0045808 |  | 0.63754 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.51934
| 1000 | 1000 | 2024-11-19 23:30 | 0.63557 | 0 | 0.5239 | 0.0045857 | 1 | #N/A | 0.51933 | 0.51931 | 0.5193 | 0.5193 | 0.54224
| 5875000 | 987 | 2024-11-19 23:31 | 0.61971 | 0 | 0.53763 | 0.0046214 |  | 0.63615 | 0.54225 | 0.54226 | 0.54226 | 0.54226 | 0.51915
| 1000 | 1000 | 2024-11-19 23:32 | 0.63554 | 0 | 0.52374 | 0.004617 | 1 | #N/A | 0.51914 | 0.51913 | 0.51912 | 0.51912 | 0.54221
| 5875000 | 987 | 2024-11-19 23:33 | 0.61971 | 0 | 0.53762 | 0.0046199 |  | 0.63614 | 0.54222 | 0.54224 | 0.54225 | 0.54226 | 0.51914
| 1000 | 1000 | 2024-11-19 23:33 | 0.63554 | 0 | 0.52374 | 0.0046197 | 1 | #N/A | 0.51913 | 0.51912 | 0.51911 | 0.51911 | 0.54222
| 5900000 | 1003 | 2024-11-19 23:34 | 0.61785 | 0 | 0.53759 | 0.0046517 |  | 0.63435 | 0.54222 | 0.54223 | 0.54225 | 0.54225 | 0.51898
| 1000 | 1000 | 2024-11-19 23:35 | 0.63552 | 0 | 0.5236 | 0.0046498 | 1 | #N/A | 0.51897 | 0.51896 | 0.51895 | 0.51894 | 0.5422
| 5900000 | 1003 | 2024-11-19 23:36 | 0.61786 | 0 | 0.53758 | 0.0046467 |  | 0.63434 | 0.54221 | 0.54222 | 0.54224 | 0.54225 | 0.519
| 1000 | 1000 | 2024-11-19 23:37 | 0.63554 | 0 | 0.52362 | 0.0046487 | 1 | #N/A | 0.51898 | 0.51897 | 0.51896 | 0.51895 | 0.54221
| 5900000 | 1003 | 2024-11-19 23:37 | 0.61786 | 0 | 0.53758 | 0.0046495 |  | 0.63435 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.51898
| 1000 | 1000 | 2024-11-19 23:38 | 0.63553 | 0 | 0.52361 | 0.004649 | 1 | #N/A | 0.51897 | 0.51896 | 0.51895 | 0.51894 | 0.5422
| 5922000 | 987 | 2024-11-19 23:39 | 0.61578 | 0 | 0.53755 | 0.0046796 |  | 0.63234 | 0.54221 | 0.54222 | 0.54224 | 0.54225 | 0.51883
| 1000 | 1000 | 2024-11-19 23:40 | 0.63549 | 0 | 0.52348 | 0.0046746 | 1 | #N/A | 0.51882 | 0.51881 | 0.5188 | 0.51878 | 0.54217
| 5934000 | 989 | 2024-11-19 23:40 | 0.61448 | 0 | 0.53752 | 0.0046916 |  | 0.63103 | 0.54219 | 0.5422 | 0.54221 | 0.54222 | 0.51875
| 1000 | 1000 | 2024-11-19 23:41 | 0.63553 | 0 | 0.52342 | 0.0046966 | 1 | #N/A | 0.51874 | 0.51873 | 0.51871 | 0.51871 | 0.5422
| 5940000 | 990 | 2024-11-19 23:42 | 0.61381 | 0 | 0.53753 | 0.0046975 |  | 0.63036 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.51874
| 1000 | 1000 | 2024-11-19 23:43 | 0.63554 | 0 | 0.52341 | 0.0046987 | 1 | #N/A | 0.51873 | 0.51872 | 0.51871 | 0.51871 | 0.54221
| 5952000 | 992 | 2024-11-19 23:43 | 0.61232 | 0 | 0.53752 | 0.0047116 |  | 0.62889 | 0.54222 | 0.54222 | 0.54223 | 0.54224 | 0.51867
| 1000 | 1000 | 2024-11-19 23:44 | 0.63553 | 0 | 0.52336 | 0.0047133 | 1 | #N/A | 0.51866 | 0.51865 | 0.51864 | 0.51863 | 0.54221
| 5964000 | 994 | 2024-11-19 23:45 | 0.61071 | 0 | 0.5375 | 0.0047262 |  | 0.62728 | 0.54222 | 0.54222 | 0.54223 | 0.54225 | 0.5186
| 1000 | 1000 | 2024-11-19 23:46 | 0.63553 | 0 | 0.52331 | 0.004726 | 1 | #N/A | 0.51859 | 0.51858 | 0.51858 | 0.51857 | 0.54221
| 5970000 | 995 | 2024-11-19 23:46 | 0.60985 | 0 | 0.5375 | 0.0047346 |  | 0.62642 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.51856
| 1000 | 1000 | 2024-11-19 23:47 | 0.63555 | 0 | 0.52328 | 0.0047348 | 1 | #N/A | 0.51856 | 0.51855 | 0.51855 | 0.51854 | 0.54222
| 5982000 | 997 | 2024-11-19 23:48 | 0.60802 | 0 | 0.5375 | 0.0047519 |  | 0.62463 | 0.54223 | 0.54224 | 0.54226 | 0.54227 | 0.51849
| 1000 | 1000 | 2024-11-19 23:49 | 0.63553 | 0 | 0.52321 | 0.0047475 | 1 | #N/A | 0.51848 | 0.51847 | 0.51846 | 0.51845 | 0.5422
| 5988000 | 998 | 2024-11-19 23:49 | 0.60705 | 0 | 0.53748 | 0.0047606 |  | 0.62366 | 0.54222 | 0.54223 | 0.54225 | 0.54226 | 0.51844
| 1000 | 1000 | 2024-11-19 23:50 | 0.63553 | 0 | 0.52317 | 0.0047594 | 1 | #N/A | 0.51843 | 0.51841 | 0.5184 | 0.51839 | 0.5422
| 6000000 | 1000 | 2024-11-19 23:51 | 0.60501 | 0 | 0.53746 | 0.0047711 |  | 0.62159 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.51837
| 1000 | 1000 | 2024-11-19 23:52 | 0.63553 | 0 | 0.52312 | 0.004771 | 1 | #N/A | 0.51836 | 0.51836 | 0.51835 | 0.51834 | 0.54221
| 6012000 | 1002 | 2024-11-19 23:52 | 0.60285 | 0 | 0.53745 | 0.0047905 |  | 0.61944 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.51828
| 1000 | 1000 | 2024-11-19 23:53 | 0.63553 | 0 | 0.52305 | 0.0047877 | 1 | #N/A | 0.51828 | 0.51827 | 0.51827 | 0.51825 | 0.5422
| 6018000 | 1003 | 2024-11-19 23:54 | 0.6017 | 0 | 0.53744 | 0.0047965 |  | 0.61829 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.51825
| 1000 | 1000 | 2024-11-19 23:54 | 0.63554 | 0 | 0.52302 | 0.0047978 | 1 | #N/A | 0.51824 | 0.51823 | 0.51822 | 0.51821 | 0.54221
| 6030000 | 1005 | 2024-11-19 23:55 | 0.59936 | 0 | 0.53743 | 0.0048038 |  | 0.6159 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.51822
| 1000 | 1000 | 2024-11-19 23:56 | 0.63556 | 0 | 0.523 | 0.0048087 | 1 | #N/A | 0.5182 | 0.51819 | 0.51819 | 0.51818 | 0.54223
| 6036000 | 1006 | 2024-11-19 23:57 | 0.59806 | 0 | 0.53743 | 0.0048241 |  | 0.61465 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.51813
| 1000 | 1000 | 2024-11-19 23:57 | 0.63553 | 0 | 0.52292 | 0.0048209 | 1 | #N/A | 0.51812 | 0.5181 | 0.51809 | 0.51808 | 0.5422
| 6048000 | 1008 | 2024-11-19 23:58 | 0.59549 | 0 | 0.5374 | 0.0048224 |  | 0.61199 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.51811
| 1000 | 1000 | 2024-11-19 23:59 | 0.63555 | 0 | 0.52292 | 0.0048265 | 1 | #N/A | 0.5181 | 0.5181 | 0.51809 | 0.51808 | 0.54223
| 6060000 | 1010 | 2024-11-20 | 0.59269 | 0 | 0.5374 | 0.0048509 |  | 0.60921 | 0.54224 | 0.54224 | 0.54225 | 0.54226 | 0.51799
| 1000 | 1000 | 2024-11-20 | 0.63555 | 0 | 0.52282 | 0.0048498 | 1 | #N/A | 0.51798 | 0.51798 | 0.51797 | 0.51796 | 0.54222
| 6066000 | 1011 | 2024-11-20 00:01 | 0.59119 | 0 | 0.53738 | 0.0048646 |  | 0.60773 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.51792
| 1000 | 1000 | 2024-11-20 00:02 | 0.63554 | 0 | 0.52276 | 0.0048642 | 1 | #N/A | 0.51791 | 0.5179 | 0.51789 | 0.51788 | 0.54222
| 6078000 | 1013 | 2024-11-20 00:03 | 0.58821 | 0 | 0.53737 | 0.0048751 |  | 0.60469 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.51787
| 1000 | 1000 | 2024-11-20 00:03 | 0.63554 | 0 | 0.52272 | 0.0048743 | 1 | #N/A | 0.51786 | 0.51785 | 0.51784 | 0.51784 | 0.54222
| 6090000 | 1015 | 2024-11-20 00:04 | 0.58503 | 0 | 0.53735 | 0.0048917 |  | 0.60148 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.51778
| 1000 | 1000 | 2024-11-20 00:05 | 0.63553 | 0 | 0.52265 | 0.0048909 | 1 | #N/A | 0.51777 | 0.51776 | 0.51775 | 0.51774 | 0.54221
| 6096000 | 1016 | 2024-11-20 00:06 | 0.58337 | 0 | 0.53733 | 0.0049 |  | 0.59981 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.51773
| 1000 | 1000 | 2024-11-20 00:06 | 0.63553 | 0 | 0.52261 | 0.0049001 | 1 | #N/A | 0.51772 | 0.51771 | 0.5177 | 0.5177 | 0.54221
| 6108000 | 1018 | 2024-11-20 00:07 | 0.57996 | 0 | 0.53732 | 0.0049142 |  | 0.59636 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.51766
| 1000 | 1000 | 2024-11-20 00:08 | 0.63553 | 0 | 0.52255 | 0.0049159 | 1 | #N/A | 0.51765 | 0.51763 | 0.51762 | 0.51761 | 0.54221
| 6114000 | 1019 | 2024-11-20 00:09 | 0.57818 | 0 | 0.53731 | 0.0049233 |  | 0.59456 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.51761
| 1000 | 1000 | 2024-11-20 00:09 | 0.63551 | 0 | 0.5225 | 0.0049216 | 1 | #N/A | 0.5176 | 0.51758 | 0.51757 | 0.51756 | 0.54219
| 6125000 | 980 | 2024-11-20 00:10 | 0.57488 | 0 | 0.53728 | 0.0049329 |  | 0.59119 | 0.5422 | 0.54221 | 0.54222 | 0.54222 | 0.51755
| 1000 | 1000 | 2024-11-20 00:11 | 0.63552 | 0 | 0.52245 | 0.0049352 | 1 | #N/A | 0.51753 | 0.51752 | 0.51751 | 0.5175 | 0.54219
| 6150000 | 984 | 2024-11-20 00:11 | 0.56676 | 0 | 0.53725 | 0.0049665 |  | 0.58295 | 0.5422 | 0.54221 | 0.54222 | 0.54223 | 0.51739
| 1000 | 1000 | 2024-11-20 00:12 | 0.63554 | 0 | 0.52233 | 0.0049703 | 1 | #N/A | 0.51738 | 0.51736 | 0.51736 | 0.51734 | 0.54221
| 6150000 | 984 | 2024-11-20 00:13 | 0.56676 | 0 | 0.53727 | 0.0049687 |  | 0.58296 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.51739
| 1000 | 1000 | 2024-11-20 00:14 | 0.63553 | 0 | 0.52234 | 0.0049669 | 1 | #N/A | 0.51738 | 0.51737 | 0.51737 | 0.51736 | 0.5422
| 6150000 | 984 | 2024-11-20 00:14 | 0.56676 | 0 | 0.53726 | 0.004967 |  | 0.58295 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.51739
| 1000 | 1000 | 2024-11-20 00:15 | 0.63555 | 0 | 0.52234 | 0.0049714 | 1 | #N/A | 0.51738 | 0.51737 | 0.51736 | 0.51735 | 0.54222
| 6175000 | 988 | 2024-11-20 00:16 | 0.55796 | 0 | 0.53724 | 0.0050013 |  | 0.57403 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.51724
| 1000 | 1000 | 2024-11-20 00:17 | 0.63553 | 0 | 0.5222 | 0.0050011 | 1 | #N/A | 0.51722 | 0.5172 | 0.51719 | 0.51718 | 0.5422
| 6175000 | 988 | 2024-11-20 00:17 | 0.55793 | 0 | 0.53722 | 0.0050043 |  | 0.574 | 0.54222 | 0.54222 | 0.54223 | 0.54224 | 0.51721
| 1000 | 1000 | 2024-11-20 00:18 | 0.63551 | 0 | 0.52218 | 0.0050002 | 1 | #N/A | 0.5172 | 0.51719 | 0.51718 | 0.51717 | 0.54218
| 6200000 | 992 | 2024-11-20 00:19 | 0.54842 | 0 | 0.53719 | 0.0050268 |  | 0.56429 | 0.5422 | 0.54221 | 0.54222 | 0.54223 | 0.51708
| 1000 | 1000 | 2024-11-20 00:20 | 0.63553 | 0 | 0.52209 | 0.0050298 | 1 | #N/A | 0.51707 | 0.51706 | 0.51705 | 0.51705 | 0.54221
| 6200000 | 992 | 2024-11-20 00:20 | 0.54846 | 0 | 0.5372 | 0.0050294 |  | 0.56434 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.51708
| 1000 | 1000 | 2024-11-20 00:21 | 0.63551 | 0 | 0.52207 | 0.0050288 | 1 | #N/A | 0.51707 | 0.51705 | 0.51704 | 0.51702 | 0.54219
| 6200000 | 992 | 2024-11-20 00:22 | 0.54844 | 0 | 0.53719 | 0.0050291 |  | 0.56432 | 0.5422 | 0.54221 | 0.54222 | 0.54224 | 0.51707
| 1000 | 1000 | 2024-11-20 00:23 | 0.63551 | 0 | 0.52207 | 0.0050304 | 1 | #N/A | 0.51706 | 0.51705 | 0.51704 | 0.51702 | 0.54219
| 6225000 | 996 | 2024-11-20 00:23 | 0.53824 | 0 | 0.53716 | 0.0050612 |  | 0.55393 | 0.5422 | 0.54222 | 0.54223 | 0.54224 | 0.51692
| 1000 | 1000 | 2024-11-20 00:24 | 0.63551 | 0 | 0.52195 | 0.0050613 | 1 | #N/A | 0.5169 | 0.51689 | 0.51688 | 0.51687 | 0.54219
| 6225000 | 996 | 2024-11-20 00:25 | 0.53823 | 0 | 0.53716 | 0.0050625 |  | 0.55392 | 0.5422 | 0.54221 | 0.54223 | 0.54224 | 0.51691
| 1000 | 1000 | 2024-11-20 00:26 | 0.63552 | 0 | 0.52194 | 0.005065 | 1 | #N/A | 0.51689 | 0.51688 | 0.51686 | 0.51685 | 0.5422
| 6250000 | 1000 | 2024-11-20 00:26 | 0.52731 | 0 | 0.53713 | 0.0050938 |  | 0.54277 | 0.54221 | 0.54222 | 0.54222 | 0.54224 | 0.51675
| 1000 | 1000 | 2024-11-20 00:27 | 0.63553 | 0 | 0.52183 | 0.0050948 | 1 | #N/A | 0.51674 | 0.51674 | 0.51672 | 0.51672 | 0.5422
| 6250000 | 1000 | 2024-11-20 00:28 | 0.52735 | 0 | 0.53715 | 0.0050879 |  | 0.54281 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.5168
| 1000 | 1000 | 2024-11-20 00:29 | 0.63554 | 0 | 0.52186 | 0.0050891 | 1 | #N/A | 0.51679 | 0.51677 | 0.51676 | 0.51675 | 0.54221
| 6250000 | 1000 | 2024-11-20 00:29 | 0.5273 | 0 | 0.53714 | 0.0050979 |  | 0.54278 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.51675
| 1000 | 1000 | 2024-11-20 00:30 | 0.63553 | 0 | 0.52181 | 0.0050998 | 1 | #N/A | 0.51673 | 0.51672 | 0.5167 | 0.51669 | 0.54221
| 6275000 | 1004 | 2024-11-20 00:31 | 0.51571 | 0 | 0.5371 | 0.0051308 |  | 0.53094 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.51658
| 1000 | 1000 | 2024-11-20 00:32 | 0.63554 | 0 | 0.52169 | 0.0051316 | 1 | #N/A | 0.51657 | 0.51656 | 0.51655 | 0.51655 | 0.54221
| 6275000 | 1004 | 2024-11-20 00:32 | 0.51569 | 0 | 0.5371 | 0.0051308 |  | 0.53093 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.51658
| 1000 | 1000 | 2024-11-20 00:33 | 0.63553 | 0 | 0.52169 | 0.0051277 | 1 | #N/A | 0.51657 | 0.51656 | 0.51655 | 0.51655 | 0.5422
| 6300000 | 1008 | 2024-11-20 00:34 | 0.50345 | 0 | 0.53706 | 0.0051628 |  | 0.51842 | 0.5422 | 0.54222 | 0.54222 | 0.54223 | 0.51641
| 1000 | 1000 | 2024-11-20 00:35 | 0.63551 | 0 | 0.52154 | 0.0051606 | 1 | #N/A | 0.51639 | 0.51639 | 0.51638 | 0.51636 | 0.54218
| 6300000 | 1008 | 2024-11-20 00:35 | 0.50343 | 0 | 0.53706 | 0.0051607 |  | 0.5184 | 0.5422 | 0.54221 | 0.54222 | 0.54223 | 0.51641
| 1000 | 1000 | 2024-11-20 00:36 | 0.63552 | 0 | 0.52154 | 0.0051627 | 1 | #N/A | 0.5164 | 0.51639 | 0.51637 | 0.51636 | 0.54219
| 6300000 | 1008 | 2024-11-20 00:37 | 0.50344 | 0 | 0.53706 | 0.0051642 |  | 0.51842 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.5164
| 1000 | 1000 | 2024-11-20 00:38 | 0.63551 | 0 | 0.52154 | 0.0051612 | 1 | #N/A | 0.5164 | 0.51638 | 0.51637 | 0.51636 | 0.54219
| 6325000 | 1012 | 2024-11-20 00:38 | 0.49053 | 0 | 0.53702 | 0.0051883 |  | 0.5052 | 0.5422 | 0.54221 | 0.54221 | 0.54222 | 0.51627
| 1000 | 1000 | 2024-11-20 00:39 | 0.63552 | 0 | 0.52143 | 0.0051908 | 1 | #N/A | 0.51626 | 0.51625 | 0.51624 | 0.51623 | 0.5422
| 6325000 | 1012 | 2024-11-20 00:40 | 0.49052 | 0 | 0.53704 | 0.0051894 |  | 0.50519 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.51628
| 1000 | 1000 | 2024-11-20 00:40 | 0.63552 | 0 | 0.52144 | 0.0051887 | 1 | #N/A | 0.51627 | 0.51626 | 0.51625 | 0.51624 | 0.5422
| 6350000 | 1016 | 2024-11-20 00:41 | 0.477 | 0 | 0.53699 | 0.0052235 |  | 0.49136 | 0.5422 | 0.54221 | 0.54222 | 0.54223 | 0.5161
| 1000 | 1000 | 2024-11-20 00:42 | 0.63553 | 0 | 0.5213 | 0.0052246 | 1 | #N/A | 0.51609 | 0.51609 | 0.51607 | 0.51606 | 0.5422
| 6360000 | 1007 | 2024-11-20 00:43 | 0.4714 | 0 | 0.537 | 0.0052359 |  | 0.48564 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.51605
| 1000 | 1000 | 2024-11-20 00:43 | 0.63554 | 0 | 0.52126 | 0.0052365 | 1 | #N/A | 0.51604 | 0.51603 | 0.51601 | 0.51601 | 0.5422
| 6360000 | 1007 | 2024-11-20 00:44 | 0.47143 | 0 | 0.53699 | 0.0052345 |  | 0.48565 | 0.54222 | 0.54222 | 0.54223 | 0.54224 | 0.51605
| 1000 | 1000 | 2024-11-20 00:45 | 0.63553 | 0 | 0.52126 | 0.0052354 | 1 | #N/A | 0.51604 | 0.51603 | 0.51603 | 0.51602 | 0.54221
| 6375000 | 1003 | 2024-11-20 00:46 | 0.46289 | 0 | 0.53698 | 0.0052632 |  | 0.47695 | 0.54222 | 0.54224 | 0.54225 | 0.54225 | 0.51593
| 1000 | 1000 | 2024-11-20 00:46 | 0.63552 | 0 | 0.52115 | 0.0052601 | 1 | #N/A | 0.51591 | 0.5159 | 0.51589 | 0.51588 | 0.54219
| 6375000 | 1003 | 2024-11-20 00:47 | 0.46287 | 0 | 0.53697 | 0.0052583 |  | 0.47691 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.51593
| 1000 | 1000 | 2024-11-20 00:48 | 0.63552 | 0 | 0.52116 | 0.0052581 | 1 | #N/A | 0.51592 | 0.51591 | 0.5159 | 0.51588 | 0.54219
| 6400000 | 992 | 2024-11-20 00:49 | 0.44823 | 0 | 0.53695 | 0.0052841 |  | 0.46189 | 0.54221 | 0.54223 | 0.54224 | 0.54225 | 0.51581
| 1000 | 1000 | 2024-11-20 00:49 | 0.63554 | 0 | 0.52107 | 0.0052859 | 1 | #N/A | 0.5158 | 0.51579 | 0.51578 | 0.51577 | 0.54221
| 6400000 | 992 | 2024-11-20 00:50 | 0.44824 | 0 | 0.53695 | 0.0052893 |  | 0.46191 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.5158
| 1000 | 1000 | 2024-11-20 00:51 | 0.63553 | 0 | 0.52106 | 0.0052889 | 1 | #N/A | 0.51578 | 0.51577 | 0.51576 | 0.51575 | 0.54221
| 6400000 | 992 | 2024-11-20 00:52 | 0.44825 | 0 | 0.53694 | 0.0052874 |  | 0.46191 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.51579
| 1000 | 1000 | 2024-11-20 00:52 | 0.63555 | 0 | 0.52106 | 0.005289 | 1 | #N/A | 0.51579 | 0.51578 | 0.51577 | 0.51576 | 0.54222
| 6400000 | 992 | 2024-11-20 00:53 | 0.44826 | 0 | 0.53696 | 0.0052839 |  | 0.46192 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.51583
| 1000 | 1000 | 2024-11-20 00:54 | 0.63555 | 0 | 0.52108 | 0.0052854 | 1 | #N/A | 0.51581 | 0.5158 | 0.51579 | 0.51578 | 0.54222
| 6450000 | 989 | 2024-11-20 00:55 | 0.41739 | 0 | 0.53688 | 0.0053552 |  | 0.43029 | 0.54223 | 0.54223 | 0.54224 | 0.54225 | 0.51546
| 1000 | 1000 | 2024-11-20 00:55 | 0.63554 | 0 | 0.52079 | 0.0053557 | 1 | #N/A | 0.51545 | 0.51544 | 0.51543 | 0.51542 | 0.54221
| 6450000 | 989 | 2024-11-20 00:56 | 0.41739 | 0 | 0.53689 | 0.005354 |  | 0.43028 | 0.54222 | 0.54224 | 0.54225 | 0.54225 | 0.51547
| 1000 | 1000 | 2024-11-20 00:57 | 0.63553 | 0 | 0.5208 | 0.0053518 | 1 | #N/A | 0.51546 | 0.51545 | 0.51545 | 0.51543 | 0.54221
| 6450000 | 989 | 2024-11-20 00:58 | 0.41737 | 0 | 0.53688 | 0.0053538 |  | 0.43026 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.51546
| 1000 | 1000 | 2024-11-20 00:58 | 0.63554 | 0 | 0.52079 | 0.0053551 | 1 | #N/A | 0.51545 | 0.51544 | 0.51543 | 0.51542 | 0.54221
| 6450000 | 989 | 2024-11-20 00:59 | 0.41735 | 0 | 0.53688 | 0.0053567 |  | 0.43024 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.51545
| 1000 | 1000 | 2024-11-20 01:00 | 0.63553 | 0 | 0.52079 | 0.0053554 | 1 | #N/A | 0.51545 | 0.51543 | 0.51543 | 0.51542 | 0.54221
| 6450000 | 989 | 2024-11-20 01:01 | 0.41739 | 0 | 0.53688 | 0.005354 |  | 0.43028 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.51546
| 1000 | 1000 | 2024-11-20 01:01 | 0.63556 | 0 | 0.52079 | 0.0053577 | 1 | #N/A | 0.51545 | 0.51544 | 0.51543 | 0.51542 | 0.54222
| 6500000 | 1001 | 2024-11-20 01:02 | 0.38486 | 0 | 0.53682 | 0.005424 |  | 0.39692 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.51512
| 1000 | 1000 | 2024-11-20 01:03 | 0.63552 | 0 | 0.52051 | 0.0054198 | 1 | #N/A | 0.51511 | 0.51509 | 0.51509 | 0.51508 | 0.54219
| 6500000 | 1001 | 2024-11-20 01:04 | 0.38483 | 0 | 0.53681 | 0.0054196 |  | 0.39687 | 0.54221 | 0.54223 | 0.54224 | 0.54225 | 0.51514
| 1000 | 1000 | 2024-11-20 01:04 | 0.63554 | 0 | 0.52053 | 0.0054223 | 1 | #N/A | 0.51513 | 0.51512 | 0.5151 | 0.51509 | 0.54222
| 6500000 | 1001 | 2024-11-20 01:05 | 0.38484 | 0 | 0.53681 | 0.0054204 |  | 0.39689 | 0.54223 | 0.54223 | 0.54224 | 0.54224 | 0.51513
| 1000 | 1000 | 2024-11-20 01:06 | 0.63551 | 0 | 0.52052 | 0.0054185 | 1 | #N/A | 0.51512 | 0.5151 | 0.51509 | 0.51508 | 0.54219
| 6500000 | 1001 | 2024-11-20 01:06 | 0.38483 | 0 | 0.53679 | 0.0054195 |  | 0.39687 | 0.5422 | 0.54221 | 0.54222 | 0.54222 | 0.51511
| 1000 | 1000 | 2024-11-20 01:07 | 0.63553 | 0 | 0.52051 | 0.0054224 | 1 | #N/A | 0.5151 | 0.51509 | 0.51508 | 0.51507 | 0.54219
| 6500000 | 1001 | 2024-11-20 01:08 | 0.38486 | 0 | 0.53681 | 0.0054199 |  | 0.3969 | 0.54221 | 0.54222 | 0.54224 | 0.54226 | 0.51513
| 1000 | 1000 | 2024-11-20 01:09 | 0.63553 | 0 | 0.52052 | 0.0054205 | 1 | #N/A | 0.51512 | 0.51511 | 0.5151 | 0.51509 | 0.54221
| 6540000 | 981 | 2024-11-20 01:09 | 0.35784 | 0 | 0.53676 | 0.0054656 |  | 0.36913 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.5149
| 1000 | 1000 | 2024-11-20 01:10 | 0.63556 | 0 | 0.52034 | 0.0054714 | 1 | #N/A | 0.51488 | 0.51486 | 0.51486 | 0.51486 | 0.54222
| 6540000 | 981 | 2024-11-20 01:11 | 0.35784 | 0 | 0.53678 | 0.0054727 |  | 0.36915 | 0.54223 | 0.54225 | 0.54226 | 0.54226 | 0.51489
| 1000 | 1000 | 2024-11-20 01:12 | 0.63554 | 0 | 0.52033 | 0.0054704 | 1 | #N/A | 0.51487 | 0.51486 | 0.51485 | 0.51484 | 0.54221
| 6560000 | 984 | 2024-11-20 01:12 | 0.34407 | 0 | 0.53674 | 0.0054932 |  | 0.35498 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.51477
| 1000 | 1000 | 2024-11-20 01:13 | 0.63555 | 0 | 0.52024 | 0.0054953 | 1 | #N/A | 0.51475 | 0.51475 | 0.51474 | 0.51473 | 0.54222
| 6560000 | 984 | 2024-11-20 01:14 | 0.34411 | 0 | 0.53675 | 0.0054954 |  | 0.35503 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.51477
| 1000 | 1000 | 2024-11-20 01:15 | 0.63554 | 0 | 0.52024 | 0.0054949 | 1 | #N/A | 0.51476 | 0.51475 | 0.51474 | 0.51473 | 0.54222
| 6580000 | 987 | 2024-11-20 01:15 | 0.33022 | 0 | 0.53673 | 0.0055135 |  | 0.34073 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.51468
| 1000 | 1000 | 2024-11-20 01:16 | 0.63557 | 0 | 0.52018 | 0.0055151 | 1 | #N/A | 0.51467 | 0.51466 | 0.51466 | 0.51465 | 0.54224
| 6580000 | 987 | 2024-11-20 01:17 | 0.33022 | 0 | 0.53674 | 0.0055129 |  | 0.34072 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.51469
| 1000 | 1000 | 2024-11-20 01:18 | 0.63558 | 0 | 0.52018 | 0.0055162 | 1 | #N/A | 0.51468 | 0.51467 | 0.51467 | 0.51466 | 0.54225
| 6600000 | 1001 | 2024-11-20 01:18 | 0.31625 | 0 | 0.53672 | 0.0055471 |  | 0.32638 | 0.54226 | 0.54226 | 0.54227 | 0.54228 | 0.51453
| 1000 | 1000 | 2024-11-20 01:19 | 0.63558 | 0 | 0.52006 | 0.0055461 | 1 | #N/A | 0.51453 | 0.51452 | 0.51451 | 0.51451 | 0.54225
| 6600000 | 1001 | 2024-11-20 01:20 | 0.31621 | 0 | 0.53671 | 0.0055464 |  | 0.32634 | 0.54225 | 0.54226 | 0.54226 | 0.54227 | 0.51453
| 1000 | 1000 | 2024-11-20 01:21 | 0.63558 | 0 | 0.52005 | 0.005547 | 1 | #N/A | 0.51452 | 0.51451 | 0.5145 | 0.51449 | 0.54224
| 6600000 | 1001 | 2024-11-20 01:21 | 0.31621 | 0 | 0.53672 | 0.0055458 |  | 0.32634 | 0.54225 | 0.54226 | 0.54227 | 0.54227 | 0.51453
| 1000 | 1000 | 2024-11-20 01:22 | 0.63558 | 0 | 0.52005 | 0.0055465 | 1 | #N/A | 0.51452 | 0.51451 | 0.51451 | 0.51449 | 0.54224
| 6620000 | 993 | 2024-11-20 01:23 | 0.30212 | 0 | 0.53668 | 0.0055791 |  | 0.31187 | 0.54225 | 0.54225 | 0.54226 | 0.54226 | 0.51436
| 1000 | 1000 | 2024-11-20 01:24 | 0.63556 | 0 | 0.5199 | 0.0055798 | 1 | #N/A | 0.51434 | 0.51433 | 0.51431 | 0.51431 | 0.54222
| 6625000 | 1007 | 2024-11-20 01:24 | 0.29857 | 0 | 0.53666 | 0.00559 |  | 0.30822 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.5143
| 1000 | 1000 | 2024-11-20 01:25 | 0.63554 | 0 | 0.51986 | 0.0055877 | 1 | #N/A | 0.51429 | 0.51428 | 0.51427 | 0.51426 | 0.54221
| 6640000 | 996 | 2024-11-20 01:26 | 0.28801 | 0 | 0.53664 | 0.0056036 |  | 0.29733 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.51422
| 1000 | 1000 | 2024-11-20 01:27 | 0.63555 | 0 | 0.51981 | 0.0056025 | 1 | #N/A | 0.51422 | 0.51421 | 0.5142 | 0.51419 | 0.54222
| 6660000 | 999 | 2024-11-20 01:27 | 0.2739 | 0 | 0.53661 | 0.0056337 |  | 0.28282 | 0.54223 | 0.54223 | 0.54224 | 0.54226 | 0.51407
| 1000 | 1000 | 2024-11-20 01:28 | 0.63553 | 0 | 0.51967 | 0.0056332 | 1 | #N/A | 0.51406 | 0.51405 | 0.51403 | 0.51402 | 0.54221
| 6660000 | 999 | 2024-11-20 01:29 | 0.27386 | 0 | 0.53659 | 0.005636 |  | 0.28278 | 0.54221 | 0.54223 | 0.54223 | 0.54224 | 0.51405
| 1000 | 1000 | 2024-11-20 01:30 | 0.63551 | 0 | 0.51965 | 0.0056348 | 1 | #N/A | 0.51403 | 0.51402 | 0.51401 | 0.514 | 0.54219
| 6660000 | 999 | 2024-11-20 01:30 | 0.27386 | 0 | 0.53659 | 0.0056309 |  | 0.28277 | 0.5422 | 0.54221 | 0.54222 | 0.54223 | 0.51406
| 1000 | 1000 | 2024-11-20 01:31 | 0.63553 | 0 | 0.51966 | 0.005634 | 1 | #N/A | 0.51405 | 0.51403 | 0.51403 | 0.51402 | 0.5422
| 6680000 | 1002 | 2024-11-20 01:32 | 0.25977 | 0 | 0.53657 | 0.0056589 |  | 0.26827 | 0.54221 | 0.54222 | 0.54224 | 0.54225 | 0.51393
| 1000 | 1000 | 2024-11-20 01:32 | 0.63553 | 0 | 0.51957 | 0.0056593 | 1 | #N/A | 0.51392 | 0.51391 | 0.51391 | 0.51391 | 0.54221
| 6680000 | 1002 | 2024-11-20 01:33 | 0.25975 | 0 | 0.53657 | 0.0056584 |  | 0.26825 | 0.54222 | 0.54222 | 0.54224 | 0.54225 | 0.51394
| 1000 | 1000 | 2024-11-20 01:34 | 0.63554 | 0 | 0.51958 | 0.0056578 | 1 | #N/A | 0.51393 | 0.51393 | 0.51392 | 0.51391 | 0.54221
| 6700000 | 1005 | 2024-11-20 01:35 | 0.2457 | 0 | 0.53655 | 0.0056865 |  | 0.25378 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.5138
| 1000 | 1000 | 2024-11-20 01:35 | 0.63555 | 0 | 0.51946 | 0.005689 | 1 | #N/A | 0.51379 | 0.51378 | 0.51377 | 0.51376 | 0.54222
| 6700000 | 1005 | 2024-11-20 01:36 | 0.2457 | 0 | 0.53656 | 0.0056872 |  | 0.25378 | 0.54223 | 0.54224 | 0.54225 | 0.54227 | 0.51381
| 1000 | 1000 | 2024-11-20 01:37 | 0.63555 | 0 | 0.51947 | 0.0056873 | 1 | #N/A | 0.5138 | 0.51379 | 0.51378 | 0.51377 | 0.54222
| 6720000 | 1008 | 2024-11-20 01:38 | 0.23172 | 0 | 0.53653 | 0.0057118 |  | 0.23937 | 0.54223 | 0.54225 | 0.54225 | 0.54226 | 0.51369
| 1000 | 1000 | 2024-11-20 01:38 | 0.63556 | 0 | 0.51938 | 0.0057119 | 1 | #N/A | 0.51368 | 0.51367 | 0.51367 | 0.51366 | 0.54223
| 6720000 | 1008 | 2024-11-20 01:39 | 0.23173 | 0 | 0.53653 | 0.0057102 |  | 0.23938 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.51369
| 1000 | 1000 | 2024-11-20 01:40 | 0.63555 | 0 | 0.51939 | 0.005709 | 1 | #N/A | 0.51369 | 0.51368 | 0.51367 | 0.51366 | 0.54222
| 6750000 | 999 | 2024-11-20 01:41 | 0.21091 | 0 | 0.53649 | 0.0057621 |  | 0.21794 | 0.54224 | 0.54225 | 0.54225 | 0.54226 | 0.51344
| 1000 | 1000 | 2024-11-20 01:41 | 0.63553 | 0 | 0.51917 | 0.0057574 | 1 | #N/A | 0.51343 | 0.51342 | 0.51341 | 0.51339 | 0.5422
| 6750000 | 999 | 2024-11-20 01:42 | 0.2109 | 0 | 0.53647 | 0.0057606 |  | 0.21794 | 0.54221 | 0.54223 | 0.54224 | 0.54224 | 0.51343
| 1000 | 1000 | 2024-11-20 01:43 | 0.6355 | 0 | 0.51916 | 0.0057565 | 1 | #N/A | 0.51342 | 0.5134 | 0.51339 | 0.51338 | 0.54218
| 6750000 | 999 | 2024-11-20 01:44 | 0.2109 | 0 | 0.53646 | 0.0057599 |  | 0.21793 | 0.5422 | 0.54221 | 0.54223 | 0.54224 | 0.51342
| 1000 | 1000 | 2024-11-20 01:44 | 0.63552 | 0 | 0.51915 | 0.0057612 | 1 | #N/A | 0.5134 | 0.51339 | 0.51339 | 0.51338 | 0.5422
| 6760000 | 1014 | 2024-11-20 01:45 | 0.20407 | 0 | 0.53646 | 0.0057615 |  | 0.21088 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.51341
| 1000 | 1000 | 2024-11-20 01:46 | 0.63553 | 0 | 0.51915 | 0.0057642 | 1 | #N/A | 0.5134 | 0.51339 | 0.51338 | 0.51337 | 0.5422
| 6780000 | 1017 | 2024-11-20 01:47 | 0.19047 | 0 | 0.53643 | 0.0057951 |  | 0.19686 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.51325
| 1000 | 1000 | 2024-11-20 01:47 | 0.63553 | 0 | 0.51902 | 0.0057956 | 1 | #N/A | 0.51325 | 0.51323 | 0.51322 | 0.5132 | 0.5422
| 6800000 | 986 | 2024-11-20 01:48 | 0.17702 | 0 | 0.53641 | 0.0058172 |  | 0.18298 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.51315
| 1000 | 1000 | 2024-11-20 01:49 | 0.63555 | 0 | 0.51894 | 0.0058191 | 1 | #N/A | 0.51314 | 0.51312 | 0.51311 | 0.51311 | 0.54222
| 6800000 | 986 | 2024-11-20 01:50 | 0.17699 | 0 | 0.53642 | 0.0058217 |  | 0.18296 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.51314
| 1000 | 1000 | 2024-11-20 01:50 | 0.63555 | 0 | 0.51893 | 0.0058228 | 1 | #N/A | 0.51312 | 0.51311 | 0.5131 | 0.51309 | 0.54222
| 6825000 | 1001 | 2024-11-20 01:51 | 0.16052 | 0 | 0.53638 | 0.0058592 |  | 0.16596 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.51295
| 1000 | 1000 | 2024-11-20 01:52 | 0.63551 | 0 | 0.51877 | 0.0058531 | 1 | #N/A | 0.51294 | 0.51293 | 0.51291 | 0.5129 | 0.54219
| 6825000 | 1001 | 2024-11-20 01:52 | 0.1605 | 0 | 0.53636 | 0.0058558 |  | 0.16595 | 0.5422 | 0.54221 | 0.54222 | 0.54224 | 0.51294
| 1000 | 1000 | 2024-11-20 01:53 | 0.6355 | 0 | 0.51876 | 0.0058549 | 1 | #N/A | 0.51293 | 0.51291 | 0.5129 | 0.51288 | 0.54218
| 6825000 | 1001 | 2024-11-20 01:54 | 0.16052 | 0 | 0.53636 | 0.0058542 |  | 0.16597 | 0.54219 | 0.54221 | 0.54222 | 0.54222 | 0.51294
| 1000 | 1000 | 2024-11-20 01:55 | 0.6355 | 0 | 0.51876 | 0.0058552 | 1 | #N/A | 0.51292 | 0.51291 | 0.5129 | 0.51289 | 0.54218
| 6825000 | 1001 | 2024-11-20 01:55 | 0.16052 | 0 | 0.53636 | 0.0058569 |  | 0.16597 | 0.5422 | 0.54221 | 0.54223 | 0.54224 | 0.51293
| 1000 | 1000 | 2024-11-20 01:56 | 0.63551 | 0 | 0.51876 | 0.0058571 | 1 | #N/A | 0.51292 | 0.51291 | 0.5129 | 0.5129 | 0.54219
| 6825000 | 1001 | 2024-11-20 01:57 | 0.16052 | 0 | 0.53636 | 0.0058554 |  | 0.16596 | 0.54221 | 0.54221 | 0.54222 | 0.54223 | 0.51294
| 1000 | 1000 | 2024-11-20 01:58 | 0.6355 | 0 | 0.51876 | 0.005855 | 1 | #N/A | 0.51293 | 0.51291 | 0.5129 | 0.51289 | 0.54218
| 6875000 | 1001 | 2024-11-20 01:58 | 0.12871 | 0 | 0.53629 | 0.0059197 |  | 0.13313 | 0.54219 | 0.5422 | 0.54221 | 0.54222 | 0.51261
| 1000 | 1000 | 2024-11-20 01:59 | 0.63553 | 0 | 0.51851 | 0.0059241 | 1 | #N/A | 0.5126 | 0.51259 | 0.51258 | 0.51257 | 0.5422
| 6875000 | 1001 | 2024-11-20 02:00 | 0.12872 | 0 | 0.53631 | 0.0059207 |  | 0.13313 | 0.54222 | 0.54223 | 0.54224 | 0.54224 | 0.51263
| 1000 | 1000 | 2024-11-20 02:01 | 0.63552 | 0 | 0.51852 | 0.0059187 | 1 | #N/A | 0.51262 | 0.51261 | 0.5126 | 0.51259 | 0.5422
| 6875000 | 1001 | 2024-11-20 02:01 | 0.12874 | 0 | 0.5363 | 0.0059195 |  | 0.13316 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.51263
| 1000 | 1000 | 2024-11-20 02:02 | 0.63554 | 0 | 0.51853 | 0.0059229 | 1 | #N/A | 0.51262 | 0.51261 | 0.5126 | 0.51259 | 0.54222
| 6900000 | 989 | 2024-11-20 02:03 | 0.11351 | 0 | 0.53629 | 0.0059536 |  | 0.11742 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.51247
| 1000 | 1000 | 2024-11-20 02:04 | 0.63554 | 0 | 0.5184 | 0.0059527 | 1 | #N/A | 0.51246 | 0.51245 | 0.51244 | 0.51243 | 0.54221
| 6900000 | 989 | 2024-11-20 02:04 | 0.11348 | 0 | 0.53628 | 0.0059531 |  | 0.11739 | 0.54221 | 0.54222 | 0.54224 | 0.54225 | 0.51246
| 1000 | 1000 | 2024-11-20 02:05 | 0.63552 | 0 | 0.51839 | 0.0059522 | 1 | #N/A | 0.51245 | 0.51244 | 0.51244 | 0.51243 | 0.5422
| 6900000 | 989 | 2024-11-20 02:06 | 0.11349 | 0 | 0.53627 | 0.0059513 |  | 0.11741 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.51247
| 1000 | 1000 | 2024-11-20 02:07 | 0.63553 | 0 | 0.51839 | 0.0059526 | 1 | #N/A | 0.51246 | 0.51245 | 0.51244 | 0.51242 | 0.5422
| 6900000 | 989 | 2024-11-20 02:07 | 0.11348 | 0 | 0.53628 | 0.0059527 |  | 0.11739 | 0.54222 | 0.54222 | 0.54223 | 0.54224 | 0.51247
| 1000 | 1000 | 2024-11-20 02:08 | 0.63554 | 0 | 0.5184 | 0.005954 | 1 | #N/A | 0.51246 | 0.51245 | 0.51244 | 0.51243 | 0.54221
| 6900000 | 989 | 2024-11-20 02:09 | 0.1135 | 0 | 0.53628 | 0.0059542 |  | 0.11742 | 0.54223 | 0.54223 | 0.54224 | 0.54225 | 0.51247
| 1000 | 1000 | 2024-11-20 02:10 | 0.63554 | 0 | 0.5184 | 0.0059519 | 1 | #N/A | 0.51246 | 0.51245 | 0.51245 | 0.51244 | 0.54221
| 6960000 | 986 | 2024-11-20 02:10 | 0.079045 | 0 | 0.5362 | 0.0060258 |  | 0.081805 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.5121
| 1000 | 1000 | 2024-11-20 02:11 | 0.63556 | 0 | 0.51811 | 0.0060284 | 1 | #N/A | 0.5121 | 0.51209 | 0.51208 | 0.51206 | 0.54223
| 6960000 | 986 | 2024-11-20 02:12 | 0.079056 | 0 | 0.53622 | 0.0060268 |  | 0.081819 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.51211
| 1000 | 1000 | 2024-11-20 02:13 | 0.63556 | 0 | 0.51811 | 0.0060278 | 1 | #N/A | 0.5121 | 0.51209 | 0.51208 | 0.51208 | 0.54223
| 6960000 | 986 | 2024-11-20 02:13 | 0.079032 | 0 | 0.53622 | 0.006026 |  | 0.081792 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.51212
| 1000 | 1000 | 2024-11-20 02:14 | 0.63557 | 0 | 0.51812 | 0.0060276 | 1 | #N/A | 0.5121 | 0.5121 | 0.51209 | 0.51209 | 0.54223
| 6960000 | 986 | 2024-11-20 02:15 | 0.07905 | 0 | 0.53622 | 0.0060264 |  | 0.081813 | 0.54224 | 0.54225 | 0.54225 | 0.54225 | 0.51212
| 1000 | 1000 | 2024-11-20 02:16 | 0.63557 | 0 | 0.51812 | 0.006031 | 1 | #N/A | 0.5121 | 0.51209 | 0.51208 | 0.51207 | 0.54224
| 7000000 | 1001 | 2024-11-20 02:16 | 0.057916 | 0 | 0.53617 | 0.0060857 |  | 0.059961 | 0.54225 | 0.54225 | 0.54226 | 0.54226 | 0.51182
| 1000 | 1000 | 2024-11-20 02:17 | 0.63554 | 0 | 0.51788 | 0.0060841 | 1 | #N/A | 0.51181 | 0.5118 | 0.51179 | 0.51178 | 0.54221
| 7000000 | 1001 | 2024-11-20 02:18 | 0.057912 | 0 | 0.53616 | 0.0060788 |  | 0.059954 | 0.54223 | 0.54223 | 0.54224 | 0.54224 | 0.51184
| 1000 | 1000 | 2024-11-20 02:19 | 0.63554 | 0 | 0.51789 | 0.0060799 | 1 | #N/A | 0.51183 | 0.51182 | 0.5118 | 0.51179 | 0.54221
| 7000000 | 1001 | 2024-11-20 02:19 | 0.057909 | 0 | 0.53615 | 0.0060827 |  | 0.059952 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.51181
| 1000 | 1000 | 2024-11-20 02:20 | 0.63554 | 0 | 0.51788 | 0.0060833 | 1 | #N/A | 0.51181 | 0.5118 | 0.51179 | 0.51178 | 0.54221
| 7000000 | 1001 | 2024-11-20 02:21 | 0.057918 | 0 | 0.53615 | 0.0060835 |  | 0.059962 | 0.54223 | 0.54223 | 0.54224 | 0.54225 | 0.51182
| 1000 | 1000 | 2024-11-20 02:22 | 0.63553 | 0 | 0.51788 | 0.0060823 | 1 | #N/A | 0.51181 | 0.5118 | 0.51179 | 0.51178 | 0.54221
| 7000000 | 1001 | 2024-11-20 02:22 | 0.057905 | 0 | 0.53615 | 0.0060841 |  | 0.059949 | 0.54222 | 0.54223 | 0.54225 | 0.54226 | 0.51182
| 1000 | 1000 | 2024-11-20 02:23 | 0.63554 | 0 | 0.51788 | 0.0060843 | 1 | #N/A | 0.51181 | 0.5118 | 0.51179 | 0.51179 | 0.54222
| 7050000 | 987 | 2024-11-20 02:24 | 0.033814 | 0 | 0.53609 | 0.0061579 |  | 0.035023 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.51145
| 1000 | 1000 | 2024-11-20 02:24 | 0.63552 | 0 | 0.51758 | 0.0061534 | 1 | #N/A | 0.51144 | 0.51144 | 0.51142 | 0.51141 | 0.54219
| 7050000 | 987 | 2024-11-20 02:25 | 0.033799 | 0 | 0.53607 | 0.0061532 |  | 0.035006 | 0.5422 | 0.54221 | 0.54222 | 0.54224 | 0.51145
| 1000 | 1000 | 2024-11-20 02:26 | 0.63552 | 0 | 0.51758 | 0.0061534 | 1 | #N/A | 0.51144 | 0.51143 | 0.51142 | 0.51141 | 0.54219
| 7050000 | 987 | 2024-11-20 02:27 | 0.033827 | 0 | 0.53607 | 0.0061559 |  | 0.035036 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.51144
| 1000 | 1000 | 2024-11-20 02:27 | 0.63551 | 0 | 0.51757 | 0.0061543 | 1 | #N/A | 0.51143 | 0.51142 | 0.51141 | 0.5114 | 0.54218
| 7050000 | 987 | 2024-11-20 02:28 | 0.033807 | 0 | 0.53606 | 0.0061537 |  | 0.035014 | 0.5422 | 0.54221 | 0.54222 | 0.54223 | 0.51145
| 1000 | 1000 | 2024-11-20 02:29 | 0.63552 | 0 | 0.51757 | 0.0061548 | 1 | #N/A | 0.51143 | 0.51142 | 0.51141 | 0.5114 | 0.54219
| 7080000 | 1003 | 2024-11-20 02:30 | 0.020634 | 0 | 0.53604 | 0.0061849 |  | 0.021374 | 0.5422 | 0.54222 | 0.54223 | 0.54224 | 0.5113
| 1000 | 1000 | 2024-11-20 02:30 | 0.63555 | 0 | 0.51746 | 0.0061891 | 1 | #N/A | 0.51129 | 0.51128 | 0.51127 | 0.51126 | 0.54222
| 7080000 | 1003 | 2024-11-20 02:31 | 0.020626 | 0 | 0.53606 | 0.0061914 |  | 0.021368 | 0.54223 | 0.54224 | 0.54225 | 0.54226 | 0.51129
| 1000 | 1000 | 2024-11-20 02:32 | 0.63553 | 0 | 0.51745 | 0.0061889 | 1 | #N/A | 0.51128 | 0.51126 | 0.51126 | 0.51125 | 0.54221
| 7080000 | 1003 | 2024-11-20 02:33 | 0.020611 | 0 | 0.53604 | 0.0061855 |  | 0.021351 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.5113
| 1000 | 1000 | 2024-11-20 02:33 | 0.63555 | 0 | 0.51746 | 0.006189 | 1 | #N/A | 0.51129 | 0.51127 | 0.51127 | 0.51126 | 0.54222
| 7100000 | 994 | 2024-11-20 02:34 | 0.012434 | 0 | 0.53602 | 0.0062195 |  | 0.012883 | 0.54223 | 0.54223 | 0.54225 | 0.54225 | 0.51114
| 1000 | 1000 | 2024-11-20 02:35 | 0.63553 | 0 | 0.51733 | 0.0062172 | 1 | #N/A | 0.51113 | 0.51112 | 0.51111 | 0.5111 | 0.5422
| 7100000 | 994 | 2024-11-20 02:36 | 0.012447 | 0 | 0.53601 | 0.0062188 |  | 0.012897 | 0.54221 | 0.54222 | 0.54224 | 0.54225 | 0.51114
| 1000 | 1000 | 2024-11-20 02:36 | 0.63551 | 0 | 0.51732 | 0.0062163 | 1 | #N/A | 0.51112 | 0.51111 | 0.5111 | 0.51109 | 0.54219
| 7125000 | 1007 | 2024-11-20 02:37 | 0.0035247 | 0 | 0.53597 | 0.0062475 |  | 0.0036526 | 0.5422 | 0.54221 | 0.54223 | 0.54224 | 0.51098
| 1000 | 1000 | 2024-11-20 02:38 | 0.63553 | 0 | 0.5172 | 0.0062497 | 1 | #N/A | 0.51097 | 0.51095 | 0.51095 | 0.51094 | 0.5422
| 7125000 | 1007 | 2024-11-20 02:39 | 0.0035393 | 0 | 0.53597 | 0.0062486 |  | 0.0036678 | 0.5422 | 0.54221 | 0.54223 | 0.54224 | 0.51098
| 1000 | 1000 | 2024-11-20 02:39 | 0.63553 | 0 | 0.5172 | 0.0062505 | 1 | #N/A | 0.51097 | 0.51096 | 0.51095 | 0.51095 | 0.54221
| 7150000 | 1001 | 2024-11-20 02:40 | 0.0069192 | 0 | 0.53596 | 0.0062792 |  | 0.0071718 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.51084
| 1000 | 1000 | 2024-11-20 02:41 | 0.63553 | 0 | 0.51708 | 0.0062793 | 1 | #N/A | 0.51082 | 0.51081 | 0.5108 | 0.51079 | 0.5422
| 7150000 | 1001 | 2024-11-20 02:42 | 0.0069631 | 0 | 0.53594 | 0.0062807 |  | 0.0072172 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.51082
| 1000 | 1000 | 2024-11-20 02:42 | 0.63553 | 0 | 0.51707 | 0.0062817 | 1 | #N/A | 0.5108 | 0.5108 | 0.51079 | 0.51078 | 0.5422
| 7150000 | 1001 | 2024-11-20 02:43 | 0.0068717 | 0 | 0.53595 | 0.0062769 |  | 0.0071223 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.51084
| 1000 | 1000 | 2024-11-20 02:44 | 0.63553 | 0 | 0.51709 | 0.0062797 | 1 | #N/A | 0.51083 | 0.51081 | 0.5108 | 0.51079 | 0.54221
| 7150000 | 1001 | 2024-11-20 02:44 | 0.0069367 | 0 | 0.53595 | 0.0062802 |  | 0.0071897 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.51083
| 1000 | 1000 | 2024-11-20 02:45 | 0.63554 | 0 | 0.51709 | 0.0062824 | 1 | #N/A | 0.51082 | 0.51081 | 0.5108 | 0.51079 | 0.54222
| 7200000 | 996 | 2024-11-20 02:46 | 0.022608 | 0 | 0.53589 | 0.0063428 |  | 0.023442 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.51052
| 1000 | 1000 | 2024-11-20 02:47 | 0.63554 | 0 | 0.51684 | 0.0063439 | 1 | #N/A | 0.51051 | 0.5105 | 0.51049 | 0.51048 | 0.54221
| 7200000 | 996 | 2024-11-20 02:47 | 0.02258 | 0 | 0.53589 | 0.0063436 |  | 0.023413 | 0.54222 | 0.54223 | 0.54223 | 0.54224 | 0.51051
| 1000 | 1000 | 2024-11-20 02:48 | 0.63553 | 0 | 0.51683 | 0.0063419 | 1 | #N/A | 0.51051 | 0.5105 | 0.51049 | 0.51048 | 0.5422
| 7200000 | 996 | 2024-11-20 02:49 | 0.0226 | 0 | 0.53588 | 0.0063429 |  | 0.023434 | 0.54221 | 0.54222 | 0.54223 | 0.54223 | 0.51051
| 1000 | 1000 | 2024-11-20 02:50 | 0.63552 | 0 | 0.51682 | 0.0063445 | 1 | #N/A | 0.51049 | 0.51049 | 0.51047 | 0.51046 | 0.5422
| 7200000 | 996 | 2024-11-20 02:50 | 0.022603 | 0 | 0.53588 | 0.0063435 |  | 0.023436 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.51051
| 1000 | 1000 | 2024-11-20 02:51 | 0.63553 | 0 | 0.51683 | 0.0063431 | 1 | #N/A | 0.5105 | 0.51049 | 0.51048 | 0.51047 | 0.5422
| 7200000 | 996 | 2024-11-20 02:52 | 0.022574 | 0 | 0.53589 | 0.0063424 |  | 0.023406 | 0.54221 | 0.54223 | 0.54224 | 0.54225 | 0.51052
| 1000 | 1000 | 2024-11-20 02:53 | 0.63553 | 0 | 0.51683 | 0.0063423 | 1 | #N/A | 0.51051 | 0.5105 | 0.51049 | 0.51048 | 0.5422
| 7250000 | 986 | 2024-11-20 02:53 | 0.035745 | 0 | 0.53582 | 0.0064148 |  | 0.037079 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.51016
| 1000 | 1000 | 2024-11-20 02:54 | 0.63552 | 0 | 0.51654 | 0.0064143 | 1 | #N/A | 0.51015 | 0.51013 | 0.51012 | 0.51011 | 0.5422
| 7250000 | 986 | 2024-11-20 02:55 | 0.035743 | 0 | 0.53581 | 0.0064135 |  | 0.037076 | 0.54221 | 0.54222 | 0.54223 | 0.54223 | 0.51015
| 1000 | 1000 | 2024-11-20 02:56 | 0.63552 | 0 | 0.51654 | 0.0064131 | 1 | #N/A | 0.51014 | 0.51013 | 0.51012 | 0.51011 | 0.54219
| 7250000 | 986 | 2024-11-20 02:56 | 0.035752 | 0 | 0.53581 | 0.0064145 |  | 0.037086 | 0.54221 | 0.54222 | 0.54224 | 0.54225 | 0.51016
| 1000 | 1000 | 2024-11-20 02:57 | 0.63551 | 0 | 0.51654 | 0.0064138 | 1 | #N/A | 0.51014 | 0.51013 | 0.51012 | 0.5101 | 0.54219
| 7250000 | 986 | 2024-11-20 02:58 | 0.035757 | 0 | 0.53581 | 0.0064173 |  | 0.037092 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.51014
| 1000 | 1000 | 2024-11-20 02:59 | 0.63551 | 0 | 0.51653 | 0.0064127 | 1 | #N/A | 0.51013 | 0.51012 | 0.51012 | 0.51011 | 0.54218
| 7250000 | 986 | 2024-11-20 02:59 | 0.035737 | 0 | 0.5358 | 0.0064142 |  | 0.037069 | 0.5422 | 0.54221 | 0.54223 | 0.54223 | 0.51015
| 1000 | 1000 | 2024-11-20 03:00 | 0.63553 | 0 | 0.51655 | 0.0064143 | 1 | #N/A | 0.51014 | 0.51014 | 0.51013 | 0.51012 | 0.5422
| 7290000 | 972 | 2024-11-20 03:01 | 0.044368 | 0 | 0.53577 | 0.0064645 |  | 0.046037 | 0.54222 | 0.54224 | 0.54224 | 0.54225 | 0.50991
| 1000 | 1000 | 2024-11-20 03:02 | 0.63552 | 0 | 0.51635 | 0.0064625 | 1 | #N/A | 0.5099 | 0.50989 | 0.50988 | 0.50986 | 0.5422
| 7300000 | 1022 | 2024-11-20 03:02 | 0.046274 | 0 | 0.53575 | 0.0064773 |  | 0.048018 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.50984
| 1000 | 1000 | 2024-11-20 03:03 | 0.63553 | 0 | 0.51629 | 0.0064808 | 1 | #N/A | 0.50982 | 0.50981 | 0.5098 | 0.50979 | 0.54221
| 7300000 | 1022 | 2024-11-20 03:04 | 0.046304 | 0 | 0.53576 | 0.0064809 |  | 0.048051 | 0.54222 | 0.54223 | 0.54224 | 0.54225 | 0.50983
| 1000 | 1000 | 2024-11-20 03:05 | 0.63551 | 0 | 0.51628 | 0.0064751 | 1 | #N/A | 0.50982 | 0.50982 | 0.50981 | 0.5098 | 0.54219
| 7300000 | 1022 | 2024-11-20 03:05 | 0.046265 | 0 | 0.53573 | 0.0064792 |  | 0.048009 | 0.5422 | 0.54221 | 0.54222 | 0.54223 | 0.50982
| 1000 | 1000 | 2024-11-20 03:06 | 0.63551 | 0 | 0.51627 | 0.0064789 | 1 | #N/A | 0.5098 | 0.50979 | 0.50979 | 0.50978 | 0.54219
| 7320000 | 976 | 2024-11-20 03:07 | 0.049747 | 0 | 0.53572 | 0.0064902 |  | 0.051624 | 0.5422 | 0.54221 | 0.54222 | 0.54222 | 0.50976
| 1000 | 1000 | 2024-11-20 03:08 | 0.63555 | 0 | 0.51624 | 0.0064946 | 1 | #N/A | 0.50975 | 0.50974 | 0.50975 | 0.50974 | 0.54222
| 7335000 | 978 | 2024-11-20 03:08 | 0.052112 | 0 | 0.53572 | 0.00652 |  | 0.05409 | 0.54223 | 0.54224 | 0.54224 | 0.54225 | 0.50964
| 1000 | 1000 | 2024-11-20 03:09 | 0.63553 | 0 | 0.51613 | 0.006519 | 1 | #N/A | 0.50963 | 0.50962 | 0.5096 | 0.50959 | 0.5422
| 7350000 | 980 | 2024-11-20 03:10 | 0.054247 | 0 | 0.53568 | 0.0065442 |  | 0.056315 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.50951
| 1000 | 1000 | 2024-11-20 03:11 | 0.63552 | 0 | 0.51601 | 0.0065444 | 1 | #N/A | 0.50949 | 0.50948 | 0.50946 | 0.50945 | 0.54219
| 7350000 | 980 | 2024-11-20 03:11 | 0.054228 | 0 | 0.53568 | 0.0065445 |  | 0.056294 | 0.5422 | 0.54221 | 0.54223 | 0.54224 | 0.5095
| 1000 | 1000 | 2024-11-20 03:12 | 0.63552 | 0 | 0.51601 | 0.0065451 | 1 | #N/A | 0.50949 | 0.50947 | 0.50946 | 0.50945 | 0.54219
| 7375000 | 1003 | 2024-11-20 03:13 | 0.057323 | 0 | 0.53565 | 0.0065714 |  | 0.059516 | 0.5422 | 0.54222 | 0.54223 | 0.54223 | 0.50936
| 1000 | 1000 | 2024-11-20 03:14 | 0.63552 | 0 | 0.51591 | 0.0065716 | 1 | #N/A | 0.50936 | 0.50935 | 0.50933 | 0.50932 | 0.5422
| 7375000 | 1003 | 2024-11-20 03:14 | 0.057297 | 0 | 0.53565 | 0.0065739 |  | 0.059491 | 0.5422 | 0.54221 | 0.54222 | 0.54223 | 0.50935
| 1000 | 1000 | 2024-11-20 03:15 | 0.63553 | 0 | 0.5159 | 0.0065752 | 1 | #N/A | 0.50934 | 0.50933 | 0.50931 | 0.50931 | 0.5422
| 7375000 | 1003 | 2024-11-20 03:16 | 0.057313 | 0 | 0.53565 | 0.006573 |  | 0.059507 | 0.54221 | 0.54222 | 0.54222 | 0.54223 | 0.50935
| 1000 | 1000 | 2024-11-20 03:16 | 0.63552 | 0 | 0.5159 | 0.0065738 | 1 | #N/A | 0.50934 | 0.50933 | 0.50932 | 0.50932 | 0.5422
| 7400000 | 999 | 2024-11-20 03:17 | 0.059857 | 0 | 0.53562 | 0.0066018 |  | 0.062158 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.50922
| 1000 | 1000 | 2024-11-20 03:18 | 0.63555 | 0 | 0.51579 | 0.0066061 | 1 | #N/A | 0.5092 | 0.50919 | 0.50918 | 0.50917 | 0.54222
| 7400000 | 999 | 2024-11-20 03:19 | 0.05986 | 0 | 0.53564 | 0.0066035 |  | 0.062163 | 0.54223 | 0.54224 | 0.54225 | 0.54225 | 0.50923
| 1000 | 1000 | 2024-11-20 03:20 | 0.63555 | 0 | 0.5158 | 0.0066038 | 1 | #N/A | 0.50921 | 0.5092 | 0.50919 | 0.50918 | 0.54222
| 7400000 | 999 | 2024-11-20 03:20 | 0.059853 | 0 | 0.53563 | 0.0066059 |  | 0.062155 | 0.54222 | 0.54223 | 0.54225 | 0.54226 | 0.50921
| 1000 | 1000 | 2024-11-20 03:21 | 0.63553 | 0 | 0.51579 | 0.0066038 | 1 | #N/A | 0.50919 | 0.50919 | 0.50919 | 0.50918 | 0.54221
| 7425000 | 990 | 2024-11-20 03:22 | 0.061817 | 0 | 0.53559 | 0.0066348 |  | 0.064207 | 0.54222 | 0.54222 | 0.54224 | 0.54224 | 0.50905
| 1000 | 1000 | 2024-11-20 03:23 | 0.63553 | 0 | 0.51566 | 0.0066354 | 1 | #N/A | 0.50904 | 0.50903 | 0.50902 | 0.50901 | 0.5422
| 7425000 | 990 | 2024-11-20 03:23 | 0.06182 | 0 | 0.53559 | 0.0066351 |  | 0.064209 | 0.54221 | 0.54222 | 0.54223 | 0.54224 | 0.50905
| 1000 | 1000 | 2024-11-20 03:24 | 0.63553 | 0 | 0.51566 | 0.0066347 | 1 | #N/A | 0.50904 | 0.50903 | 0.50902 | 0.50901 | 0.5422
| 7440000 | 992 | 2024-11-20 03:25 | 0.062766 | 0 | 0.53558 | 0.0066502 |  | 0.065198 | 0.54221 | 0.54222 | 0.54224 | 0.54225 | 0.50898
| 1000 | 1000 | 2024-11-20 03:26 | 0.63552 | 0 | 0.5156 | 0.006647 | 1 | #N/A | 0.50897 | 0.50896 | 0.50895 | 0.50894 | 0.54219
| 7455000 | 994 | 2024-11-20 03:26 | 0.063495 | 0 | 0.53555 | 0.0066681 |  | 0.065961 | 0.5422 | 0.54221 | 0.54223 | 0.54224 | 0.50888
| 1000 | 1000 | 2024-11-20 03:27 | 0.63553 | 0 | 0.51553 | 0.0066684 | 1 | #N/A | 0.50888 | 0.50887 | 0.50886 | 0.50885 | 0.5422
| 7455000 | 994 | 2024-11-20 03:28 | 0.063517 | 0 | 0.53555 | 0.0066703 |  | 0.065986 | 0.54222 | 0.54222 | 0.54223 | 0.54224 | 0.50887
| 1000 | 1000 | 2024-11-20 03:29 | 0.63553 | 0 | 0.51552 | 0.0066711 | 1 | #N/A | 0.50886 | 0.50885 | 0.50884 | 0.50884 | 0.5422
| 7470000 | 996 | 2024-11-20 03:29 | 0.064095 | 0 | 0.53553 | 0.0066928 |  | 0.066595 | 0.54221 | 0.54221 | 0.54222 | 0.54223 | 0.50875
| 1000 | 1000 | 2024-11-20 03:30 | 0.63552 | 0 | 0.51542 | 0.0066935 | 1 | #N/A | 0.50874 | 0.50873 | 0.50873 | 0.50872 | 0.5422
| 7485000 | 998 | 2024-11-20 03:31 | 0.064506 | 0 | 0.5355 | 0.0067137 |  | 0.06703 | 0.54221 | 0.54221 | 0.54222 | 0.54222 | 0.50865
| 1000 | 1000 | 2024-11-20 03:31 | 0.63552 | 0 | 0.51534 | 0.0067153 | 1 | #N/A | 0.50864 | 0.50863 | 0.50862 | 0.50861 | 0.5422
| 7485000 | 998 | 2024-11-20 03:32 | 0.064516 | 0 | 0.5355 | 0.006715 |  | 0.067042 | 0.5422 | 0.54221 | 0.54222 | 0.54223 | 0.50864
| 1000 | 1000 | 2024-11-20 03:33 | 0.63551 | 0 | 0.51532 | 0.0067144 | 1 | #N/A | 0.50863 | 0.50861 | 0.50861 | 0.5086 | 0.54218
| 7500000 | 1000 | [6 empty] | #DIV/0!
STOP | 1000 | 1000 | [5 empty] | 1 | #DIV/0!
Note on D2: labina: Real sine amplitude, i.e. A*sin(w*t). It is NOT rms or peak-to-peak value.
Note on D7: labina: The algorithm must generate output quantity 'rms'. That is e.g. TWM-WFFT or TWM-WRMS